## Gas Std
I. APH Method Calibration Standard Information and Input - Nominal Concentration-Based
Calibration Control Number: | APH | 10250201 |  | APH Working Gas Standards Preparation |  |  |  | LOW | HIGH
Calibration Date: | 2002-10-25
Analyst: | DGM |  |  | Summa Canister Volume (liters) |  |  |  | 6 | 6
Input Retention Times (minutes) and Calibration Stock Standard Concentration* (mg/m3) for: | 1. APH Target Analytes |  |  | Flow Rate (cc/min) |  |  |  | 25 | 400
| 2. APH Range Coponents | [6 empty] | 14.5 | 23
| 3. Internal Standards (ISs), and |  |  | Final Canister Pressure (psig) |  |  |  | 30 | 30
| 4. Tuning Standard |  |  | DILUTION FACTOR |  |  |  | 50.3308 | 1.9831
Compound | MW | Retention Time Minutes | Range | Gas Standard ACTUAL Concentration |  | Working Standard Concentration (Low) |  | Working Standard Concentration (High)
(√) APH TARGET Analyte |  |  |  | mg/m3 | PPMv | ug/m3 | PPBv | ug/m3 | PPBv
Unadjusted C5-C8 Aliphatic Hydrocarbons | N/A | N/A |  | 14 | 3.5804 | 300 | N/A | 7000 | N/A
Unadjusted C9-C12 Aliphatic Hydrocarbons | N/A | N/A |  | 5 | 0.83672 | 100 | N/A | 2500 | N/A
1,3-Butadiene ( C ) | 60.14 | 5.26 |  | 1 | 0.40655 | 19.8686 | 8.0776 | 500 | 203.2757
Isopentane | 72.15 | 6.13 | C5-C8-Al | 1 | 0.33888 | 19.8686 | 6.733 | 500 | 169.4387
Methyl-tert-butylether ( √ ) | 88.15 | 7.57 |  | 1 | 0.27737 | 19.8686 | 5.5109 | 500 | 138.6841
n-Hexane | 86.15 | 8.29 | C5-C8-Al | 1 | 0.28381 | 19.8686 | 5.6388 | 500 | 141.9037
Benzene ( √ ) | 78.11 | 9.65 |  | 1 | 0.31302 | 19.8686 | 6.2193 | 500 | 156.51
Cyclohexane | 84.16 | 9.93 | C5-C8-Al | 1 | 0.29052 | 19.8686 | 5.7722 | 500 | 145.259
2,3-Dimethylpentane | 100.21 | 10.01 | C5-C8-Al | 1 | 0.24399 | 19.8686 | 4.8477 | 500 | 121.9938
n-Heptane | 101.2 | 10.78 | C5-C8-Al | 1 | 0.2416 | 19.8686 | 4.8003 | 500 | 120.8004
Toluene ( √ ) | 92.14 | 12.72 |  | 1 | 0.26536 | 19.8686 | 5.2723 | 500 | 132.6785
n-Octane | 114.23 | 13.98 | C5-C8-Al | 1 | 0.21404 | 19.8686 | 4.2527 | 500 | 107.0209
Ethylbenzene ( √ ) | 106.17 | 16.1 |  | 1 | 0.23029 | 19.8686 | 4.5756 | 500 | 115.1455
2,3-Dimethylheptane | 128.26 | 16.16 | C9-C12-Al | 1 | 0.19063 | 19.8686 | 3.7875 | 500 | 95.3142
m- & p- Xylenes ( √ ) | 106.17 | 16.45 |  | 2 | 0.46058 | 39.7371 | 9.1511 | 1000 | 230.291
o-Xylene ( √ ) | 106.17 | 17.37 |  | 1 | 0.23029 | 19.8686 | 4.5756 | 500 | 115.1455
n-Nonane | 128.26 | 17.75 | C9-C12-Al | 1 | 0.19063 | 19.8686 | 3.7875 | 500 | 95.3142
Isopropylbenzene | 120.2 | 18.62 | C9-C10-Ar | 1 | 0.20341 | 19.8686 | 4.0415 | 500 | 101.7055
1-Methyl-3-ethylbenzene | 120.2 | 20.22 | C9-C10-Ar | 1 | 0.20341 | 19.8686 | 4.0415 | 500 | 101.7055
1,3,5-Trimethylbenzene | 120.2 | 20.43 | C9-C10-Ar | 1 | 0.20341 | 19.8686 | 4.0415 | 500 | 101.7055
n-Decane | 142.29 | 21.9 | C9-C12-Al | 1 | 0.17183 | 19.8686 | 3.4141 | 500 | 85.9161
1,2,3-Trimethylbenzene | 120.2 | 22.83 | C9-C10-Ar | 1 | 0.20341 | 19.8686 | 4.0415 | 500 | 101.7055
p-Isopropyltoluene | 134.22 | 22.83 | C9-C10-Ar | 1 | 0.18216 | 19.8686 | 3.6193 | 500 | 91.0818
Butylcyclohexane | 140.27 | 23.61 | C9-C12-Al | 1 | 0.17431 | 19.8686 | 3.4632 | 500 | 87.1533
n-Undecane | 156.31 | 26.16 | C9-C12-Al | 1 | 0.15642 | 19.8686 | 3.1078 | 500 | 78.21
n-Dodecane | 170.34 | 29.62 | C9-C12-Al | 1 | 0.14354 | 19.8686 | 2.8519 | 500 | 71.7682
Naphthalene ( √ ) | 128.16 | 29.27 |  | 1 | 0.19078 | 19.8686 | 3.7905 | 500 | 95.3886
C5-C8 Aliphatic Hydrocarbons | N/A | N/A |  | 6 | N/A | 120 | N/A | 3000 | N/A
C9-C12 Aliphatic Hydrocarbons | N/A | N/A |  | 6 | N/A | 120 | N/A | 3000 | N/A
C9-C10 Aromatic Hydrocarbons | N/A | N/A |  | 5 | N/A | 100 | N/A | 2500 | N/A
Total Reported Hydrocarbons | N/A | N/A |  | N/A | N/A | 500 | N/A | 13000 | N/A
1 Working Standard - Concentration of APH Gas Standard prepared from APH Stock Standard introduced into gas-sampling system
Internal Standards (IS) & BFB Tuning Data | MW | Retention Time Minutes | IS Gas Conc PPBv | IS Gas Std Conc ug/m3 | IS Flow Rate cc/min | IS Flow Duration minutes | On-Column IS(x) Mass (ηg) | Nominal IS Gas Conc ug/m3
Bromochloromethane, IS(1) | 129.43 | 9.65 | 100 | 529.3661 | 20 | 1 | 10.5873 | 42.3493
1,4-Difluorobenzene, IS(2) | 114.09 | 9.95 | 100 | 466.6258 | 20 | 1 | 9.3325 | 37.3301
Chlorobenzene-d5, IS(3) | 117.56 | 15.4 | 100 | 480.818 | 20 | 1 | 9.6164 | 38.4654
4-Bromofluorobenzene (BFB) | 175.01 | 18.32 | 100 | 715.7873 | 20 | 1 | 14.3157 | 57.263
II. Area Count Input for APH Calibration and Calculations of Relative Response Factors - Nominal Concentration-Based
1. INSERT Area Counts for Individual compounds for each Calibration Level
2. INSERT calibration flow rate (cc/min) and duration (min), used to establish individual Calibration Levels
Calibration Standard Level |  | 1 | 2 | 3 | 4 | 5 | 6
Working Standard (ug/m3) |  | 19.8686 | 19.8686 | 19.8686 | 500 | 500 | 500
Calibration Flow Rate (cc/min) |  | 5 | 10 | 25 | 5 | 25 | 25
Calibration Flow Duration (minutes) |  | 5 | 5 | 5 | 5 | 5 | 10
On-Column Volume (cc) |  | 25 | 50 | 125 | 25 | 125 | 250
On-Column Mass (ng) |  | 0.49671 | 0.99343 | 2.4836 | 12.5 | 62.5 | 125
Nominal* Concentration (ug/m3) |  | 1.9869 | 3.9737 | 9.9343 | 50 | 250 | 500
1. Nominal sample volume 0.25 liters
Compound | QUANT ION | AREA COUNT for Individual Standard Levels
1,3-Butadiene | 54 | 6337 | 14664 | 36660 | 183300 | 980573 | 1854649
Bromochloromethane, IS(1) | 121 | 35531 | 36214 | 39788 | 41232 | 40515 | 60078
Isopentane | 43 | 2223 | 7456 | 18640 | 93200 | 459196 | 1031595
Methyl-tert-butylether | 73 | 6747 | 17575 | 43937 | 219688 | 1424596 | 3645385
n-Hexane | 57 | 1391 | 3384 | 8460 | 42300 | 255744 | 761077
Benzene | 78 | 3556 | 8201 | 20502 | 102512 | 647506 | 1691722
1,4-Difluorobenzene, IS(2) | 114 | 143419 | 148189 | 146799 | 162114 | 172980 | 246302
Cyclohexane | 56 | 1788 | 5250 | 13125 | 65625 | 374747 | 967332
2,3-Dimethylpentane | 56 | 2999 | 6325 | 15812 | 79062 | 471028 | 1248544
n-Heptane | 43 | 1827 | 6002 | 15005 | 75025 | 489349 | 1288707
Toluene | 91 | 6092 | 12494 | 31235 | 156175 | 1086109 | 2437046
Chlorobenzene-d5, IS(3) | 117 | 316020 | 320770 | 297404 | 313031 | 336207 | 378992
n-Octane | 43 | 7869 | 15973 | 39932 | 199662 | 806445 | 1881104
Ethylbenzene | 91 | 16511 | 35273 | 88182 | 440912 | 2411322 | 4659148
2,3-Dimethylheptane | 43 | 9786 | 21985 | 54962 | 274812 | 1301031 | 2402318
m- & p- Xylenes | 91 | 26176 | 60190 | 150475 | 752375 | 3999652 | 7788338
o-Xylene | 91 | 12988 | 28608 | 71520 | 357600 | 2102284 | 4197368
n-Nonane | 43 | 6763 | 14581 | 36452 | 182262 | 1060313 | 2016617
Isopropylbenzene (120 a.m.u) | 120 | 12267 | 29472 | 73680 | 368400 | 2221982 | 3497442
Isopropylbenzene (134 a.m.u.) | 134 | 0 | 0 | 0 | 0 | 0 | 0
1-Methyl-3-ethylbenzene (120 a.m.u) | 120 | 10155 | 24432 | 61080 | 305400 | 2021310 | 4120625
1-Methyl-3-ethylbenzene (134 a.m.u.) | 134 | 0 | 0 | 0 | 0 | 0 | 0
1,3,5-Trimethylbenzene (120 a.m.u) | 120 | 9935 | 25420 | 63550 | 317750 | 1965146 | 4005654
1,3,5-Trimethylbenzene (134 a.m.u.) | 134 | 0 | 0 | 0 | 0 | 0 | 0
n-Decane | 57 | 4417 | 10654 | 26635 | 133175 | 821407 | 1690516
1,2,3-Trimethylbenzene (120 a.m.u) | 120 | 9383 | 23154 | 57885 | 289425 | 1785313 | 3709878
1,2,3-Trimethylbenzene (134 a.m.u.) | 134 | 0 | 0 | 0 | 0 | 0 | 0
p-Isopropyltoluene (120 a.m.u) | 120 | 3600.8 | 8982 | 22454.8 | 112274.8 | 737858.8 | 1607608.8
p-Isopropyltoluene (134 a.m.u.) | 134 | 9002 | 22455 | 56137 | 280687 | 1844647 | 4019022
Butylcyclohexane | 83 | 4510 | 11261 | 28152 | 140762 | 841313 | 1747933
n-Undecane | 57 | 3083 | 7327 | 18317 | 91587 | 641647 | 1558980
n-Dodecane | 57 | 3737 | 7084 | 17710 | 88550 | 666786 | 1473457
4-Bromofluorobenzene (BFB) | TUNE | 342161 | 346944 | 362579 | 373001 | 397666 | 422781
Naphthalene3 | 128 | 7042 | 10026 | 25065 | 125325 | 814108 | 1760433
C5-C8 - AlHC-(Area Correction) | N/A
C9-C12 - AlHC - (Area Correction) | N/A
C9-C10 ArHC-120 a.m.u - Area Correction | N/A
C9-C10 ArHC-134 a.m.u - Area Correction | N/A | [6 empty] | Number of Range Standards
C5-C8 Aliphatic Hydrocarbons (TI-6) | N/A | 18097 | 44390 | 110974 | 554874 | 2856509 | 7178359 | 6
C9-C12 Aliphatic Hydrocarbons (TI-6) | N/A | 32296 | 72892 | 182228 | 911148 | 5332497 | 10889821 | 6
C9-C10 Aromatic Hydrocarbons (120-5) | N/A | 45340.8 | 111460 | 278649.8 | 1393249.8 | 8731609.8 | 16941207.8 | N/A
C9-C10 Aromatic Hydrocarbons (134-5) | N/A | 9002 | 22455 | 56137 | 280687 | 1844647 | 4019022 | N/A
C9-C10 Aromatic Hydrocarbons (Total-5) | N/A | 54342.8 | 133915 | 334786.8 | 1673936.8 | 10576256.8 | 20960229.8 | 5
On-Column Volume - Volume of calibration gas (CC) introduced into gas-sampling system to establish individual Calibration Levels
III. Calculation Of Relative Response Factors for APH Analytes and Ranges - Nominal Concentration-Based
Compound | Relative Response Factors for Individual Calibration Standard Levels | [5 empty] | AVERAGE RRF | % RSD
| 1 | 2 | 3 | 4 | 5 | 6
1,3-Butadiene | 3.8015 | 4.3154 | 3.9278 | 3.7653 | 4.0999 | 2.6147 | 3.7541 | 15.8283
Isopentane | 0.29122 | 0.47266 | 0.47714 | 0.42922 | 0.39639 | 0.3127 | 0.39656 | 20.0091
Methyl-tert-butylether | 4.0475 | 5.1721 | 4.7075 | 4.5128 | 5.9564 | 5.1393 | 4.9226 | 13.3458
n-Hexane | 0.18223 | 0.21452 | 0.21656 | 0.19481 | 0.22076 | 0.2307 | 0.20993 | 8.5476
Benzene | 0.46585 | 0.51989 | 0.5248 | 0.47211 | 0.55894 | 0.5128 | 0.50907 | 6.8611
Cyclohexane | 0.23424 | 0.33282 | 0.33597 | 0.30223 | 0.32349 | 0.29322 | 0.30366 | 12.5122
2,3-Dimethylpentane | 0.39288 | 0.40097 | 0.40475 | 0.36411 | 0.4066 | 0.37846 | 0.3913 | 4.299
n-Heptane | 0.23934 | 0.38049 | 0.38409 | 0.34552 | 0.42242 | 0.39064 | 0.36042 | 17.8096
Toluene | 0.37321 | 0.37704 | 0.40666 | 0.38382 | 0.49705 | 0.49469 | 0.42208 | 13.8198
n-Octane | 1.0309 | 1.0126 | 1.0222 | 0.91953 | 0.69614 | 0.57021 | 0.87525 | 22.3746
Ethylbenzene | 1.0115 | 1.0644 | 1.1481 | 1.0836 | 1.1035 | 0.94575 | 1.0595 | 6.7577
2,3-Dimethylheptane | 0.59951 | 0.66345 | 0.71557 | 0.67538 | 0.5954 | 0.48764 | 0.62282 | 12.9613
m- & p- Xylenes | 0.80179 | 0.90819 | 0.97954 | 0.92452 | 0.9152 | 0.79047 | 0.88662 | 8.409
o-Xylene | 0.79567 | 0.86331 | 0.93114 | 0.87884 | 0.96209 | 0.85202 | 0.88051 | 6.7241
n-Nonane | 0.41431 | 0.44002 | 0.47458 | 0.44793 | 0.48524 | 0.40935 | 0.44524 | 6.9153
Isopropylbenzene | 0.7515 | 0.88939 | 0.95926 | 0.90538 | 1.0169 | 0.70994 | 0.87206 | 13.6457
1-Methyl-3-ethylbenzene | 0.62211 | 0.73729 | 0.79522 | 0.75055 | 0.92503 | 0.83644 | 0.77777 | 13.1169
1,3,5-Trimethylbenzene | 0.60864 | 0.76711 | 0.82738 | 0.78091 | 0.89933 | 0.8131 | 0.78274 | 12.3953
n-Decane | 0.27059 | 0.32151 | 0.34677 | 0.32729 | 0.37591 | 0.34315 | 0.33087 | 10.6134
1,2,3-Trimethylbenzene | 0.57482 | 0.69873 | 0.75362 | 0.71129 | 0.81703 | 0.75306 | 0.71809 | 11.3514
p-Isopropyltoluene | 0.77207 | 0.94868 | 1.0232 | 0.96575 | 1.1819 | 1.1421 | 1.0056 | 14.7082
Butylcyclohexane | 0.27629 | 0.33983 | 0.36652 | 0.34594 | 0.38502 | 0.35481 | 0.34473 | 10.783
n-Undecane | 0.18887 | 0.22111 | 0.23847 | 0.22509 | 0.29364 | 0.31645 | 0.24727 | 19.4803
n-Dodecane | 0.22894 | 0.21378 | 0.23057 | 0.21762 | 0.30515 | 0.29909 | 0.24919 | 16.6714
Naphthalene3 | 0.43141 | 0.30256 | 0.32633 | 0.308 | 0.37257 | 0.35735 | 0.3497 | 13.871
C5-C8 Aliphatic Hydrocarbons | 0.39513 | 0.46901 | 0.47344 | 0.4259 | 0.41097 | 0.36266 | 0.42285 | 10.1622
C9-C12 Aliphatic Hydrocarbons | 0.32975 | 0.36661 | 0.39541 | 0.37321 | 0.40673 | 0.36842 | 0.37336 | 7.1583
C9-C10 Aromatic Hydrocarbons | 0.66583 | 0.80824 | 0.87174 | 0.82278 | 0.96802 | 0.85094 | 0.83126 | 11.8725
Notes: | 1. Retention Time information (I), Working Standard Concentrations (I), on-Column Calibration Volumes (II) and Relative Response Factors (III) are exported to Calculation Sheet (1) as input

## LABINPUT
Sample Number: | 0162 (TM-1) |  |  |  |  | Calibration Control Number
Sample Volume (cc): | 250 |  |  |  |  | APH | 10250201
Dilution Factor: | 1
Internal Standards (IS) & Tuning Data | Quant Ion | Retention Time Minutes | On-Column IS(x) Mass ηg | Area Count | Area Count Range Corrected* | * If IS1, IS2, IS3 and BFB Area Counts (ACs) for sample are automatically removed from APH cumulative Range ACs by GC/MS data system, enter the AC removed for individual ISs and BFB in Cells 7, 8, 9 and 10, respectively, in Area Count Range Corrected column.
Bromochloromethane, IS(1) | 128 | 9.65 | 10.5873 | 31534 | 31534
1,4-Difluorobenzene, IS(2) | 114 | 9.95 | 9.3325 | 115082 | 115082
Chlorobenzene-d5, IS(3) | 117 | 15.4 | 9.6164 | 289465 | 289465
4-Bromofluorobenzene (BFB) | TUNE | 18.32 | 14.3157 | 425176 | 425176
APH Target Analytes & Range Data Calculations | Quant Ion | Retention Time | Range Start | Range End | Relative Response Factor | Area Count | Preliminary Sample Result (ηg) | Preliminary Sample Result (ug/m3)
1,3-Butadiene (√) | 54 | 5.26 | No Range |  | 3.7541 | 100452 | 8.9838 | 35.9351
Bromochloromethane, IS(1) | 128 | 9.65 | C5-C8 Aliphatic |  | N/A | 0 | N/A | N/A
Methyl-tert-butylether (√) | 73 | 7.57 | C5-C8 Aliphatic |  | 4.9226 | 162682 | 11.0957 | 44.3826
Benzene (√) | 78 | 9.65 | C5-C8 Aliphatic |  | 0.50907 | 60285 | 9.6034 | 38.4137
1,4-Difluorobenzene, IS(2) | 114 | 9.95 | C5-C8 Aliphatic |  | N/A | 0 | N/A | NA
Toluene (√) | 91 | 12.72 | C5-C8 Aliphatic |  | 0.42208 | 119314 | 9.3911 | 37.5643
Chlorobenzene-d5, IS(3) | 117 | 15.4 | C5-C8 Aliphatic |  | N/A | 0 | N/A | N/A
Ethylbenzene (√) | 91 | 16.1 | C5-C8 Aliphatic |  | 1.0595 | 325648 | 10.2111 | 40.8443
m- & p- Xylenes (√) | 91 | 16.45 | C5-C8 Aliphatic |  | 0.88662 | 518803 | 19.4393 | 77.7571
o-Xylene (√) | 91 | 17.37 | C5-C8 Aliphatic |  | 0.88051 | 250488 | 9.4508 | 37.803
4-Bromofluorobenzene (BFB) | TUNE | 18.32 | C9-C12 Aliphatic |  | N/A | 0 | N/A | N/A
Naphthalene3 (√) | 128 | 29.27 | C9-C12 Aliphatic |  | 0.3497 | 99759 | 9.477 | 37.9079
C5-C8 Aliphatic Hydrocarbons | TIC | Range | 6.03 | 17.65 | 0.42285 | 823563 | 157.9429 | 631.7717
C9-C12 Aliphatic Hydrocarbons | TIC | NA | 17.65 | 29.52 | 0.37336 | 1971741 | 175.4455 | 701.782
C9-C10 Aromatic Hydrocarbons-120 | 120 | N/A | N/A | N/A | N/A | 2256810 | NA | N/A
C9-C10 Aromatic Hydrocarbons-134 | 134 | Range | N/A | N/A | N/A | 960760 | NA | N/A
C9-C10 Aromatic Hydrocarbons-Total | Total (120 + 134) | Range | 17.38 | 29.26 | 0.83126 | 3217570 | 128.5901 | 514.3604
N/A - Not Applicable
Insert verified analytical data and sample information into appropriate YELLOW cells on the APH Sheet (1). Calculated data will automatically be inserted into appropriate BLUE cells on the sheet. The calculated results from the "PreliminarySample Result, ug/m3 column will be automatically inserted into the YELLOW input section of the APH Calculation Sheet (2). Retention Times for the APH Aliphatic Ranges will also be imported to the APH Calculation Sheet (2).

## I-Calcs
Calibration Control Number: |  |  |  | APH | #REF!
APH Range DATA
RANGE | Start Time minutes |  | End Time minutes |  | Analytical and Range Data Imported From Sheet (1)
C5-C8 Aliphatic Hydrocarbons | 6.12 |  | 17.74
C9-C12 Aliphatic Hydrocarbons | 17.74 |  | 29.61
C9-C10 Aromatic Hydrocarbons | 17.38 |  | 29.26
ANALYTICAL DATA INPUT (μg/m3):
APH Target Analytes and Ranges | MW | Eluting Range | Preliminary Concentration μg/m3 | SAMPLE NUMBER
Unadjusted C5-C8 Aliphatic Hydrocarbons | N/A | N/A | 631.7717 | 0162 (TM-1)
Unadjusted C9-C12 Aliphatic Hydrocarbons | N/A | N/A | 701.782
Unadjusted C9-C10 Aromatic Hydrocarbons | N/A | N/A | 514.3604
1,3-Butadiene | 60.14 | No Range | 35.9351
Methyl-tert-butylether | 88.15 | C5-C8-Al | 44.3826
Benzene | 78.11 | C5-C8-Al | 38.4137
Toluene | 92.14 | C5-C8-Al | 37.5643
Ethylbenzene | 106.17 | C5-C8-Al | 40.8443
m- & p- Xylenes | 106.17 | C5-C8-Al | 77.7571
o-Xylene | 106.17 | C5-C8-Al | 37.803
Naphthalene | 128.16 | C9-C12-Al | 37.9079
CALCULATIONS:
Sum of the C5-C8 Range APH Target Analytes | N/A | N/A | 276.7651
Sum of the C9-C12 Range APH Target Analytes | N/A | N/A | 37.9079
Adjusted C9-C10 Aromatic Hydrocarbon Range | N/A | N/A | 514.3604
Adjusted C9-C12 Aliphatic Hydrocarbon Range | N/A | N/A | 149.5137
Adjusted C5-C8 Aliphatic Hydrocarbon Range | N/A | N/A | 355.0066
Total Reported Hydrocarbons | N/A | N/A | 1369.4887
APH DATA REPORTING:
Target APH Analytes & Hydrocarbon Ranges | Range in which Analyte Elutes |  | Reporting Limit |  | Sample Results
|  |  | μg/m3 | ppbv | μg/m3 | ppbv
Unadjusted C5-C8 Aliphatic Hydrocarbons | N/A |  | N/A | N/A | 631.7717 | N/A
Unadjusted C9-C12 Aliphatic Hydrocarbons | N/A |  | N/A | N/A | 701.782 | N/A
1,3-Butadiene | No Range |  | 1.9869 | 0.80842 | 35.9351 | 14.6214
Methyl-tert-butylether | C5-C8-Al |  | 1.9869 | 0.55154 | 44.3826 | 12.3204
Benzene | C5-C8-Al |  | 1.9869 | 0.62243 | 38.4137 | 12.0341
Toluene | C5-C8-Al |  | 1.9869 | 0.52766 | 37.5643 | 9.9761
Ethylbenzene | C5-C8-Al |  | 1.9869 | 0.45793 | 40.8443 | 9.4138
m- & p- Xylenes | C5-C8-Al |  | 3.9737 | 0.91586 | 77.7571 | 17.9214
o-Xylene | C5-C8-Al |  | 1.9869 | 0.45793 | 37.803 | 8.7128
Naphthalene | C9-C12-Al |  | 1.9869 | 0.37936 | 37.9079 | 7.2379
C5-C8 Aliphatic Hydrocarbons | N/A |  | 11.9211 | N/A | 360 | N/A
C9-C12 Aliphatic Hydrocarbons | N/A |  | 11.9211 | N/A | 150 | N/A
C9-C10 Aromatic Hydrocarbons | N/A |  | 9.9343 | N/A | 510 | N/A
N/A - Not Applicable
Verified analytical data is directly imported from Calculation Sheet (1) and inserted into individual cells in the ANALYTICAL DATA INPUT and APH range Data sections into the Retention Time and Reported Conc. ug/m3 (YELLOW Columns), respectively. Calculation Sheet (2) will automatically insert data into Reported Conc. ppbV Column, appropriate cells in the CALCULATIONS section in the BLUE column and rounded (ROUND (num, -2)) Sample Results columns, ug/m3 and ppbV, in the APH DATA REPORTING section. Optionally, the Reporting Limit data may be inputed into the appropriate cell in the ug/m3 column (YELLOW) and the corresponding ppbV values will automatically be be inserted.