Massachusetts Department of Developmental Services Risk Management

Bed Rail Entrapment

Individuals supported by DDS who need supportive and protective devices for their beds can be at risk for injury or death if the equipment is not properly installed or is the wrong size for the bed or mattress. The following information can be helpful in preventing such an event.

What is it?

Entrapment is defined as an event in which an individual is caught, trapped or entangled in the spaces in or about the bed rail, mattress or bed frame.

What risk does it present?

Entrapment can result in serious injury or death. Deaths usually involve entrapment of the head, neck, or thorax, while most injuries involved fractures, cuts, and abrasions to the extremities.  The majority of the deaths and injuries involved elderly patients or frail adults. Individuals at high risk for entrapment include those with pre-existing conditions such as confusion, restlessness, lack of muscle control, or a combination of these factors.

How is it managed?

The FDA recommends the following actions to prevent deaths and injuries from entrapment in bed side rails:

•Inspect all bed frames, bed side rails, and mattresses as part of a regular maintenance program to identify areas of possible entrapment. Regardless of mattress width, length, and/or depth, alignment of the bed frame, bed side rail, and mattress should leave no gap wide enough to entrap a patient's head or body.

•Be aware that gaps can be created by movement or compression of the mattress which may be caused by patient weight, patient movement, or bed position.

The seven areas in the bed system where there is a potential for entrapment are identified in Figure 1.

Zone 1: Within the Rail

Zone 2: Under the Rail, Between the Rail Supports or Next to a Single Rail Support

Zone 3: Between the Rail and the Mattress

Zone 4: Under the Rail, at the Ends of the Rail

Zone 5: Between Split Bed Rail

Zone 6: Between the End of the Rail and the Side Edge of the Head or Foot Board

Zone 7: Between the Head or Foot

Figure 1

SUMMARY OF ENTRAPMENT ZONES

While recognizing that there are seven potential entrapment zones, entrapment zones 1-4 account for the majority of entrapments and deaths from entrapments. Focusing on these high risk areas allows us to prioritize strategies for deceasing risks. For dimensional limits and test methods for zones 1, 2, 3, 4, refer to FDA Guidance, Hospital Bed System Dimensional and Assessment Guidance to Reduce Entrapment (Document issued on: March 10, 2006).

Where and to whom does this apply?

All residential/day programs serving individuals that use bed rails.

Where can I find additional help?

- www.Mass.gov/DDS: http://www.mass.gov/eohhs/docs/dmr/awp/hpci-risk-bed-rail-entrapment.pdf

- NIH – www.nih.gov

- CDC – www.cdc.gov

- DPH – www.mass.gov/dph

- DMH – www.mass.gov/dmh

- http://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/GuidanceDocuments/ucm072662.htm   Hospital Bed System Dimensional and Assessment Guidance to Reduce Entrapment (Document issued on: March 10, 2006).  (document number 1537)