Current Temp

1

In the morning, view display for current temperature and check for “ALARMS”.

Daily Procedure for Monitoring Temperatures Using Fridge-Tag 2L

Repeat steps 2 & 3 for each additional day site is closed (holidays and weekends) to review min/max temperatures.

Press “READ” to view maximum (warmest) temperature from the previous day.

The Date and Time will appear under the Events column whenever “READ” button is pressed. This fulfills the requirement to monitor temperatures twice daily.

Press

“READ”

2

5

3

Press

“READ”

Press

“READ”

Press

“READ”

4

Press “READ” to view minimum (coldest) temperature from the previous day.

Press “READ” to view minimum (coldest) temperature since midnight.

Press “READ” to view the maximum (warmest) temperature since midnight.

Freezer

REFRIGERATOR

If temperature is warmer than:

-15°C (+5°F)

for >60 Minutes

If temperature is colder than:

2°C (35°F)

for >15 Minutes

If temperature is warmer than:

8°C (46°F)

for > 30 Minutes

Every AM

Immediately

Upload temp logs into MIIS and call the Vaccine Unit at 617-983-6828 if temperatures are out of range.

Office

Closed

Every PM

2-3

Repeat steps 1 through 3.

1-3