Colilert® QuickGuide

MassDEP Watershed Planning Program

CN 198.6 (updated 4/25/23)

- Check that the incubator is on and reading 35.0° C (+/-0.5°C) and record temperature on the log sheet (front of incubator). Clean all work areas with lab disinfectant. Sign chain of custody (in pocket of fridge) transferring sample custody from field crew (or fridge) to analyst.

- Allow samples to warm up to near room temperature.

- Turn on Quanti-Tray sealer to allow preheating (takes around 10 minutes). The switch is in the back near the plug. Sealer is ready to use when green light is on.

- Fill out the printed worksheet (in the large binders): batch number (next sequential number), sample IDs, lab sample #, and date/time collected.  Add lab sample ID#s for a lab duplicate (if there is enough sample to split for a lab dup), positive control (“QC”), and lab blank (“LB”).

- Line up the sample bottles and equipment. Mark the back of the Quanti-Trays with the lab ID# of each sample.

- Analysis order: work from the oldest samples to newest in the order they were collected. Sample analysis must be started within 8 hours of collection. QC samples (lab duplicates, QC samples) and dilutions can be done at any time during the analysis; lab blanks should be done last.

- Prepare the control samples for each analysis batch:

  - Blank: add sterile water to an IDEXX sample bottle to the 100-ml line.

  - Positive control: add sterile water to an IDEXX sample bottle to the 100-ml line and add a lyophilized E. coli pellet to the sample. Re-cap and mix well.

  - Lab duplicate: using 250-ml field sample, mix sample well and pour 100-ml aliquots into two new IDEXX bottles.

- SAMPLES (without dilution):  Mix the field sample (by inverting the bottle 20-25x). If necessary, pour off excess sample to the 100-ml line (bottom of meniscus). Avoid touching the inside of the bottle or lid.

Prepare any dilutions needed. (Look for any notes from the field crews indicating any suspected need for dilution.)

  - 10 ML DILUTION SAMPLES:  Pour 100 mls of sterile water into a new, sterile IDEXX vessel to as close as possible to the 100-ml line. Use a sterile pipette to pipette off 10 mls for a final volume of 90 mls of sterile water. Use the same pipette to transfer 10 mls of the original, mixed sample into the 90 mls of sterile water.

  - 100 ML DILUTION SAMPLES:  Pour 100 mls of sterile water into a new, sterile IDEXX vessel as close as possible to the 100-ml line. Then use a sterile pipette to transfer 1 ml of the original, mixed sample into the 99 mls of sterile water.  Cap and mix.

- Open Colilert reagent packet (snap off top) and add entire contents to each 100-ml sample.  Re-cap and mix thoroughly until no more reagent particles are visible. Pour into pre-labeled tray.

- Open the Quanti-Tray by pulling the foil tab away from the well side. Squeeze upper part of the tray gently so that it bends slightly to make a bigger opening.  Avoid touching the inside of the foil or tray. Pour the whole 100-ml sample into the tray. Tap the small wells to release any air bubbles.

- Seal the trays: Put the Quanti-Tray well-side down on the sealer tray. Put the tray into the sealer with beveled edge first. The sealer will automatically pull the tray in.

(NOTES:  If the Sealer stops in the middle of processing a Quanti-Tray sample, use the button with the arrows to reverse the direction of the Quanti-Tray. If the sealer does not easily accept the tray, lift the end of the tray slightly, so that the leading edge clears the opening of the sealer.)

- Incubate all prepared Colilert Quanti-Trays for 24 - 28 hours.  (If Colilert-18 is used, incubate for 18-22 hours at 35 o C.)  Record the incubation start time in the workbook.

- Thoroughly clean and disinfect the work area.

- After 24 hours (18 hours for Colilert-18), remove the trays from the incubator and record the incubation end time in the lab workbook.  Check the incubator that the temperature reading is still 35.0° C (+/-0.5°C); record temperature on the log sheet (front of incubator).

- Enterococci results: For each tray, count the number of large wells (including the biggest well) that turned yellow and number of small wells that turned yellow. (Count the well even if it has a smaller amount of sample if it is yellow.) Record the number of positive (yellow) wells in the workbook.

- E. coli results: expose each tray to 365-nm UV light by placing them into the UV light viewing cabinet. Count the number of large wells and the number of small wells fluorescing. Record in the workbook.

- Save a copy (“save as”) of the Ecoli_Lab_Book_Template from OneDrive (WPP Lab SOPs and Results 2023) with the new batch number as the file name.

- Enter all results from the printed workbook into the electronic workbook; detailed directions are available in the “Read Me” tab.  Final results will be automatically displayed in the blue-shaded “Total Coliform MPN” and “E coli MPN” columns.

- After all results have been recorded and entered into the electronic workbook, check that all required data and metadata have been recorded completely and accurately. Keep the printed worksheet in the lab binder.

NOTE: Colilert and Colilert-18 results are definitive at 24-28 and 18-22 hours, respectively.  Any positives for both total coliform and E.coli observed before the minimum time and negatives observed after the maximum time are also valid.  Positive results after the maximum time are not valid.

EXAMPLE DATA ENTRY SHEET

EXAMPLE BATCH_QC (sheet produced after data entry is compete by hitting CTR+Shift+Z)