Massachusetts Department of Environmental Protection Division of Watershed Management

STANDARD OPERATING PROCEDURE

ANalyTICAL QUANTIFICATION OF ESCHERICHIA COLI AND ENTEROCOCCI BACTERIA IN AMBIENT SURFACE WATERS Using AN ENZYME SUBSTRATE TEST (sTANDARD mETHODS 9223b)

CN 198.0

February, 2007

| Prepared by: |  | Date: |  |
| --- | --- | --- | --- |
|  | ###### Richard Chase, Quality Assurance Analyst Joan Beskenis, Environmental Analyst |  |  |
| Approved by: |  | Date: |  |
|  | ###### Arthur Johnson, Monitoring Coordinator |  |  |
| Approved by: |  | Date: |  |
|  | ###### Dennis Dunn, Program Supervisor |  |  |

### * see pdf version for valid signatures

Acknowledgements:   Several existing SOP documents for enzyme substrate procedures were used in developing this SOP, including that from the MADEP Wall Experiment Station, Massachusetts Water Resource Authority, the USEPA NE Regional Laboratory and the Ohio River Sanitation Commission.

# TABLE OF CONTENTS									Page

List of Revisions										        3

Scope and Application										        4

Summary of Methods										        4

Definitions										       	        5

Interferences											        5

Lab Safety											        5

Equipment and Supplies									        6

Sample Collection, Preservation and Storage							        7

Quality Control										 	        7

Calibration and Standardization								                     8

Detailed Procedure										        9

Preventative Maintenance									      12

Data Interpretation, Analysis and Validation							      12

Method Performance										      14

Pollution Prevention										      14

Waste Management									                   14

References											      15

Appendix A:  MPN Tables (by reference)

Appendix B:  MSDS (by reference)

Appendix C:  Sealer User’s Manual  (by reference)

Appendix D:  Quanti-Cult Procedure  (by reference)

Appendix E:  Lab Notebook Example

Appendix F:  Incubator Temperature Log

Appendix G: Waste Receipt Form

LIST OF REVISIONS

| Rev. # | Date | Description of Revision(s) | Page #s |
| --- | --- | --- | --- |
| 0 | 4/2004 | Original draft | 19 |
| 1 | 6/4/2004 | Mainly Section 10 changes, misc. minor revisions | 19 |
| 2 | 7/2006 | Minor, misc changes throughout | --- |
| 3 | 9/2006 | Added dilution requirement for marine Enterolert samples | 5, 12 |
| 4 | 2/2007 | Clarified the applicability of different Idexx reagents for waters of varying conductivities/salinities | 4-5, 10-11 |
| 5 | 3/2009 | Revised media-batch and sample batch QC | 8,9 |

1.0 	SCOPE & APPLICATION

This SOP describes the detection and enumeration of total coliform, Escherichia coli (E-coli) and Enterococci bacteria in surface water samples using the enzyme substrate test, as detailed in Standard Methods 9223B.  These methods use hydrolyzable substrates for the detection of target bacteria.

The commercially-available supplies developed by IDEXX Laboratories, Inc. were chosen to perform this procedure.   The IDEXX “Colilert” and “Colilert-18” methods are employed to detect and enumerate both total coliform and E-coli bacteria.  The IDEXX “Enterolert” method is used to detect and enumerate Enterococci (fecal streptococci) bacteria.   These defined enzyme substrate methods were recently included in an EPA Final Rule for Guidelines Establishing Test Procedures for the Analysis of Pollutants; Analytical Methods for Biological Pollutants in Ambient Water (July, 2003).

This SOP is applicable for samples collected by the DWM (or its agents) and analyzed by DWM staff at DWM’s offices/laboratory in Worcester, MA.  The main target analytes for DWM waterbody assessments, TMDL determinations and bacteria source tracking are E. coli and Enterococci (total coliform is of lesser importance).

2.0	SUMMARY OF METHODS

Because pathogenic (disease-causing) organisms are difficult to isolate and identify, fecal coliform and Enterococci bacteria are used to indicate the potential presence of pathogens in water. Coliform and fecal streptococci bacteria are widely distributed in nature and are present in the intestines of warm-blooded animals, including humans. Their presence in surface waters may indicate human or animal fecal contamination.  They can be relatively easily identified and enumerated.

Colilert (-24) and Colilert-18:

The Colilert reagent is added to a 100-mL volume of freshwater sample, the sample is poured into a multi-well tray, the tray sealed and then incubated for 24 hours (Colilert) at 35° ± 0.5°C.  The tray is then checked for color (total coliform) and fluorescent (E-coli) reactions.  The most probable number (MPN) technique utilizing a multiple well system format is used to determine the number of total coliform and E. coli per unit volume (100 mls.).   The Colilert-18 reagent provides results in only 18-22 hours.

Colilert (-24) and Colilert-18 are US EPA-approved for drinking water and source waters.  Unlike Colilert (-24) (mainly freshwater use), Colilert-18 can also be used in salt water, but for E. coli only--- not total coliform.  NOTE:  do at least a 1:10 dilution on Colilert-18 samples if they've been collected in salt/brackish water.   http://www.idexx.com/water/refs/060202711C18.pdf

The method detection limit (MDL) for total coliform and E. coli using Colilert (-24) and Colilert-18 is a MPN of 1 colony forming unit (CFU) per 100 mls.

The total coliform group is defined as all bacteria possessing the enzyme B-D-galactosidase.  The Colilert reagent contains the nutritive ortho-nitrophenyl-B-D-galactopyranoside (ONPG), which is used to detect the enzyme B-D-galactosidase of total coliforms.  The hydrolyzation of the ONPG by the enzyme produces a color change (positive=yellow) at 35.0 ° +/- 0.5° C after an incubation period of 24 to 28 hours.  The Colilert method also simultaneously detects the presence of E. coli through the hydrolyzation of the fluorogenic, nutritive substrate, 4-methyl-umbelliferyl-B-D-glucuronide (MUG) by the E. coli enzyme B-glucronidase.  This reaction produces a fluorescent product detectable when viewed under a long-wavelength (365-nm) ultraviolet (UV) light.  Non-coliform bacteria cannot metabolize the indicator nutrients.

Enterolert:

Similarly, the Enterolert reagent is added to a 100-mL volume of fresh or marine water (diluted 10X) sample, the sample is poured into a multi-well tray, the tray sealed and then incubated for 24 hours at 41° ± 0.5°C.  The tray is then checked for blue fluorescent wells indicating enterococci presence.  A MPN table is used to determine the most probable number of Enterococci per 100 mls.   The MDL for Enterococci is also a MPN of 1 CFU/100 mls.

The Enterococci, a sub-group of the fecal streptococci, contain the enzyme B-glucosidase.   The Enterolert reagent contains 4-methyl-umbelliferyl-B-D-glucoside, which reacts with the enterococci-produced enzyme B-glucosidase to produce a fluorescent blue color.

Enterolert is most applicable for marine/brackish water samples (using the required minimal 10X dilution), but can also be used for fresh water (Note:  MA. freshwater surface water quality standard is based on E. coli (Colilert), not Enterococcus). Enterolert is approved by the US EPA as a method for enterococci detection in ambient waters, including fresh, marine or estuarine surface water.  It is also an ASTM-approved method (#D 6503-99).

3.0	DEFINITIONS

RESERVED.

4.0	INTERFERENCES

The following situations represent potential interferences or complications in achieving accurate and precise results when using the Colilert, Colilert-18 and Enterolert methods.

Colilert and Colilert-18:

4.1	Non-coliform bacteria, particularly Aeromonas, and Pseudomonas species, may produce small amounts of the enzyme -D-galactosidase, but are suppressed and generally will not produce a positive response within the incubation time unless more than 104 colony-forming units (CFU)/mL are present.

4.2	Serratia species may turn the medium yellow after 24 hours of incubation but the yellow color is typically brighter than that represented by the color comparator.

4.3	Some strains of Shigella species may produce a positive fluorescence response.  This is not considered a detriment for testing the sanitary quality of water due to the pathogenic nature of Shigella.

4.4	Some water samples containing humic material may have an innate color.  If a water sample has some background color, compare inoculated Colilert sample to a control blank of the same sample.

4.5	Incubation beyond 28 hours may yield a false positive Colilert result, due to cessation of suppression of non-coliform heterotrophic bacteria.  The same is true beyond 22 hours for Colilert-18.

4.6	Presence of free chlorine in the sample may result in a transient blue color upon addition of the Colilert reagent.   To avoid this, all bacteria samples should be taken using sample bottles containing sodium thiosulfate.

4.7	Dilute brackish/marine water samples with sterile fresh water at least ten fold for Colilert-18 testing.

Enterolert:

4.8	Dilute marine water samples with sterile fresh water at least ten fold, per IDEXX Enterolert procedure.

5.0	LAB SAFETY

  - All applicable elements of DWM’s Lab Safety Plan must be adhered to at all times.

  - Samples (and positive controls) may contain organisms that are pathogenic to humans, and often the analyst works with water samples of different ranges of contamination. Handle all samples and cultures as if they are infectious.  All precautions are to be taken to minimize exposure.  These include the use personal protective equipment (lab coats, safety glasses, and protective gloves), keeping the lab work area clean and organized, working at a reasonable pace and using good judgment at all times.

  - It is recommended that all bacteria lab personnel must be immunized against the hepatitis A and B viruses.

  - All bacteria lab personnel must receive on-the-job laboratory safety training.

  - The sealer device is a burn hazard if not properly used and maintained.  See 11.1.

  - Material Safety Data Sheets (MSDS) for all chemicals used in this SOP are included in Room 229 and in Room 228.

  - Adherence to the waste storage and disposal procedures in Section 15 shall be maintained at all times at DWM and at WES.

  - Additional safety issues with regard to performance of this SOP shall be addressed by DWM’s Lab Safety Officer and/or senior staff on an as needed basis.

  - NOTE:  Pseudomonas sp. and Klebsiella pneumoniae used for media batch QC (see page 8-9) are categorized as Class 2 bio-safety hazards, whereas Enterococcus faecium and E. coli are categorized Class 1.  While Class 2 organisms can be used in this SOP, their use presents greater potential risk to analysts should an accident occur, if poor lab techniques were used or if safety protocols were not followed.

6.0	EQUIPMENT AND SUPPLIES

The following lab equipment and supplies are needed to employ the multi-well tray Colilert and Enterolert methods.  It is assumed that the necessary field supplies, such as sterile, 120 ml. bacteria sample bottles containing sodium thiosulfate, will be available for use in the field.

Colilert and Colilert-18:

  - Lab disinfectant (e.g., Roccall II 10%, Conflikt detergent) and paper towels (supply room)

  - Disposable plastic gloves and “hot” gloves (incubator and sealer) and lab coats

  - Lab notebook

  - Lab fridge and storage areas for reagents and supplies

  - Model 2X Quanti-tray sealer

  - Quanti-tray 2000 97-well sample trays  [Store in cabinet]

  - Quanti-tray 2000 tray rubber inserts (minimum 2)

  - Colilert reagent  [Store in dark at 2-30 deg. C; use within 12 months of manufacture] and Colilert-18 reagant  [Store in dark at 2-25 deg. C; use within 12 months of manufacture]

  - Sterile lab sample bottles for raw and diluted samples.  Preferably, these bottles should be high clarity (to see end points clearly), graduated for 100 mls. (accurate to within 2.5% or approx. 2-3 mls.), and be non-fluorescing.   IDEXX vessels are preferred.

  - Sterile dilution bottles (90 and 99 ml for 10X and 100X dilutions) pre-filled with sterile, buffered and/or non-buffered dilution water.  (NOTE:  IDEXX, Inc. stipulates the use of non-buffered, sterile dilution water, since their reagents are already buffered.  Use of non-buffered dilution water avoids potential analytical complications due to the presence of additional buffering agents.   In practice, a suitable buffered dilution water source may be found that does not cause analytical error or imprecision.)  [Store as directed by manufacturer]

  - Sterile pipettes (1, 10 and 50-100 ml sizes)

  - Sterile water (non-buffered).   Autoclaved sterile water available from WES in 120-250 ml. bottles.  DWM’s reagent water system (Barnstead deionization with 0.2u filtration) in Room 226 serves as a backup supply (see Table 1).  Another source may be purchased water from IDEXX.  [If sterile-bottled by WES/DWM, use within 1-2 months.   If purchased from IDEXX in sealed, 100 ml bottles, shelf life is 2 years.   Store at 2-8 deg. C or in cabinet]

  - Anti-foam solution (as needed; usu. used with P/A test)  [Store in cabinet]

  - Positive (E. coli, Enterococcus faecium) and negative (Pseudomonas, other) control cultures (for QC).  Difco Bactrol discs are acceptable. [Store at 2-8 deg. C.  Discs have 1 year shelf life]

  - Two-shelf, bench-top, mechanical-convection incubator (5-65 deg. C; +/- 1.5 deg. C uniformity; +/- 0.5 deg. C stability) with thermometer.  If also running Enterolert, two incubators (at each temperature) are preferred.  Incubators should have a temperature log sheet attached to track internal temperatures several times/day.

  - NIST-certified (or traceable) thermometer

  - 365 nm UV lamp (6 watt)

  - UV viewing cabinet

  - Colilert color comparator  NOTE:  The color comparator is the lowest color and fluorescence level at which a result can be considered positive.  A typical positive result is much more intense.   [Store in cabinet]

  - MPN tables

  - Quanti-cult QC kit   [Store at 2-8 deg. C until use; use within 18 months of manufacture]

  - Labeled, plastic bio-waste disposal bags, step-on can and labeled 32 gallon can with cover for temporary waste storage.

  - Three-clock timer

  - Autoclave (to disinfect bio-waste in bags).  This function to be performed by WES on an as-needed basis.

Enterolert:  (in addition to those listed above for Colilert)

6.25	Enterolert reagent [Store in dark at 2-30 deg. C]

Table 1:  DWM Reagent Water Quality (for backup; based on recent testing)

| Parameter | Criteria for Microbiological Analyses | Test Date | Results |
| --- | --- | --- | --- |
| Specific Conductance | < 2 uS/cm | Weekly (2004) | 1 uS/cm |
| pH | Approx. 6-7 is generally preferred | Weekly (2004) | 6-7 (for “new” DI H2O; if allowed to equilibrate to air, pH drops below 6) |
| Metals (total) for Cd, Cr, Cu, Ni, Pb and Zn | < 50 ug/l | October, 2002 (Ten Mile Watershed Metals Study) | For trip and field blanks: All DWM reagent water samples < MDL (.02-1 ug/l), except Cu which showed some “hits” from .5-4 ug/l) |
| Metals (total, combined) | < 100 ug/l | October, 2002 (Ten Mile Watershed Metals Study) | Total metals in DWM reagent water samples (mostly < MDL) were estimated to be <10 ug/l. |
| Kjeldahl Nitrogen (TKN) | < 100 ug/l | Monthly in 2003 | Invariably less than the method detection limit of 20 ug/l for ambient field blanks |
| Total Residual Chlorine | < 10 ug/l | No recent testing. (Current WES Lab MDLs (SM 4500-Cl-F) may not be able to achieve an MDL of 10 ug/l)). | DWM’s Barnstead DI water system has a dedicated carbon adsorption cartridge to adsorb chlorine, and is designed based on feedwater quality to meet Type I reagent water specifications (assuming system maintenance and feedwater consistency). |
| Fecal coliform | NA | Approx. weekly in 2003 (SM 9222D) | Rarely exceeding method detection limits (e.g., <6, <2) for ambient field blanks |
| E. coli | NA | Approx. weekly in 2003 (EPA 1603) | Rarely exceeding method detection limits (e.g., <6, <2) for ambient field blanks |
| HPC | < 100,000 CFU/100 mls. | No recent testing | --- |

7.0	SAMPLE COLLECTION, PRESERVATION AND STORAGE

For a detailed description of field sample collection for bacteria, refer to DWM Sampling Techniques for Surface Water Quality Monitoring (CN 1.21).

Briefly, new, plastic, disposable, sterile, pre-labeled, 120-mL sample bottles (usually containing a sodium thiosulfate tablet/powder for chlorine neutralization) are used to collect surface water bacteria samples.   Examples include Corning locking “flip-top” bottles (preferred) and IDEXX clear, polystyrene “vessels”.  The collection method (opening, filling and capping underwater) is consistent with the DWM SOP (CN 1.21).  After the sample is collected, a slight amount is poured off to allow at least 1/2” of headspace in the bottle.  This is required to allow for sufficient mixing of the sample at the lab prior to analysis.  Samples are placed in a clean plastic bag, which is then placed in a cooler half-filled with wet ice immediately or soon as possible after collection.  Samples (at 4 deg. C) are delivered to the lab within 6 hours from collection with the sample tracking/chain-of-custody form filled out by the collector.  At the lab, sample analysis must be initiated as soon as possible and within 8 hours of collection and two hours of receipt.

8.0	QUALITY CONTROL

Colilert and Colilert-18:

8.1	Media Lot-Specific Quality Control:  DWM’s main intent in running the following media lot QC samples is to verify that the Colilert and Colilert-18 media does not provide false positives or false negatives for E.coli.  Each lot is tested using Quanti-cult (see Appendix D) cultures or lypholyzed microorganism pellets as follows:

- 	Negative Control:  sterile water alone, AND

- Negative Control:  sterile water inoculated with non-fluorescent Pseudomonas sp. or Enterococcus faecium (i.e., total coliform & E. coli negative) or with Klebsiella pneumoniae, Enterobacter aerogenes or Enterobacter cloacae  (i.e., total coliform positive & E. coli negative), AND

- Positive Control:  sterile water inoculated with E. coli (i.e., total coliform & E. coli positive) control cultures.

Alternatively, with each new lot of media reagents, perform the following QC sample analysis with specific regard to E. coli results (using purchased QC media):

- Negative Control:  sterile water alone (lab blank), AND

- Positive Control:  sterile water spiked with E. coli., OR

- Negative Control:  sterile water spiked with an E. coli.-negative organism (see 8.1 above)

8.2	Upon receipt, check reagent packages integrity for proper seal, tears and lack of moisture.  When using, inspect appearance of Colilert reagent; it should appear dry, free-flowing and white to off-white in color.

8.3	Each lot of medium can be checked before use with the UV light to detect auto-fluorescence.  If any faint fluorescence is observed, then it should be recorded, discarded and replaced by a new lot.

8.4	If any media causes the sample to fluoresce prior to incubation, then another lot of medium should be used.

8.5	Sample Batch Specific QC:  For each lab batch, run a minimum of one lab blank and one lab duplicate (@ approx. 10% of total samples).

8.6	For reading and interpretation of results, all analysts should participate in group QC consisting of inter-analyst comparisons to confirm acceptable precision among different readers.  This is especially important for the E. coli fluorogenic end-point, for which the comparator seems to be less useful.

Enterolert:  (in addition or changed from that above for Colilert)

8.7	Media Lot-Specific Quality Control:  Similarly, run the following media lot QC samples to verify that the Enterolert media does not provide false positives or false negatives for Enterococci.  With each new lot of media purchased, run the following controls:

- 	Negative Control:  sterile water alone, AND

- Negative Control:  sterile water inoculated with non-fluorescent Serratia marcescens (gram -) and Aerococcus viridans (gram +), or E. coli (gram -), AND

- Positive Control:  sterile water inoculated with fluorescent Enterococcus faecium.

Reagent Water Testing:

8.8 	Use of DWM reagent water (Barnstead deionization system) as a backup supply for bacterial analyses (dilutions, lab blanks) requires periodic testing to ensure adequacy.  Test results are compared against criteria identified in Table 1 (exerpted from Standard Methods, Table 9020 II.

Specific conductivity and pH are measured weekly.  TKN, Fecal coliform and E. coli have been typically tested monthly.  Metals are measured as programmatic sampling dictates or annually (last tested in 11/2002; metals were below criteria).  TOC and total residual chlorine have not typically been tested for by DWM.  In general, past measurements have indicated that the DWM Barnstead deionization system, when operating normally (e.g., resistance of about 18.3, maintained correctly, etc.), may provide sufficient purity to perform bacterial analyses.

9.0	CALIBRATION AND STANDARDIZATION

  - Sample Containers:  No testing of sample containers used to measure sample volume is planned.  Containers are assumed to meet accuracy limits of < +/- 2.5%.

  - Incubator:  The incubator has an internal temperature monitoring device and maintains a temperature of 35° ± 0.5°C for E. coli.  Monitor temperatures both in air and “totally immersed” in water using total immersion incubator thermometers—should be the same.

Because pre-heated, air-type incubators may not bring water sample(s) to the specified incubation temperature of 35 °C quickly, false-negative results could result.  Therefore, the time it takes for water samples (or a set of 100-ml, water samples, depending on normal use) to reach 35 °C must be accounted for and should be known to ensure that the specified incubation period occurs (typ. >24 hrs.).  See also 10.25, 10.29 discussing additional incubation periods.

Temperature should be recorded for days in use at least twice per day with readings separated by at least 4 hours.

10.0	DETAILED PROCEDURE (assumes reagent lot has been tested and passed; see 8.1)

Colilert and Colilert-18 (and Enterolert, subject to changes/additions, as specified herein; see 10.34-37):

  - Coordinate and plan analytical work with DWM field survey staff regarding the quantity and timing of sample delivery. The maximum number of trays allowed in a DWM-Worcester incubator at one time is about 60.  For each DWM survey, one primary Colilert analyst and one backup analyst will be pre-identified and pre-scheduled to run the survey samples at DWM-Worcester.  If the primary analyst is not available for any reason, the backup analyst must be available to run them. Also schedule the Colilert equipment and incubator to make sure that it is not being used at the scheduled time by anyone else.  NOTE:   Depending on timing, available reagents and logistics, decide on which reagent will be used (Colilert (24 hour) or Colilert-18).

  - At the proper time, turn on Quanti-Tray sealer to allow preheating (usually takes around 10 minutes).  Sealer is ready to use when green light is on.

  - Prepare all work areas by cleaning with lab disinfectant.

  - Fill out lab notebook with required information, in anticipation of sample delivery.  Include and verify the date that the media lot was tested and passed.

  - Check incubator to ensure it is ON, maintaining the proper temperature and has sufficient space for sample trays when they are ready.  In order to minimize heat gain in the lab and provide adequate ventilation, make sure to turn ON the ceiling vent in Room 229 (switch is located behind Isotemp fridge in Room 227).

  - Record incubator temps to the nearest 0.5 deg. C by rounding (or 0.1 if thermometer allows).  For example, if liquid level on a 1 deg. C graduated thermometer is less than but closer to 35.0 than to the midpoint between 34 and 35, round to 35.0, not 34.5.

  - Prepare counter work area with the necessary dilution bottles, reagents, trays, pipettes, etc., leaving sufficient space to prepare samples for incubation.

  - When samples arrive, transfer custody of the samples by removing them from the plastic bag in the ice chest and checking off each sample on the COC form.  When all samples have been transferred, sign the COC form and record the time.  Review the COC form and discuss with sample delivery staff for information regarding samples that are suspected of containing a high MPN value.

  - Optional:   Depending on the amount of samples, perform sample tray preparation /incubation in groups resulting in approx. 30 minute incubation intervals.   If not being used in the current 30 minute “sub-batch”, keep samples on ice until needed.  In all cases, keep the main incubation period the same for all samples.

  - Mark the Lab Notebook with date, sample ID, lab identification number, date and time of collection, date and time of receipt, date and time of analysis, volume of original sample analyzed (dilution rate), lab QC sample information, start/end times of incubation, and the analyst’s name.   It is at this point that the dilution scheme for each sample is identified.

  - Use a Sharpie pen to mark all sample bottles with “HI” suspected of having high counts.

  - With a waterproof wax pencil (so as not to tear the paper on the back of the Quanti-Tray) gently mark the Quanti-Trays with the lab identification number of each sample.

  - For each sample (including dilutions), invert sample bottle 20-25 times to ensure complete mixing of the sample.  DO NOT TOUCH THE INSIDES OF CONTAINERS OR SAMPLE WATER.  Inspect each IDEXX vessel for inside cleanliness; if particles or other contamination is observed, place the container aside with other contaminated, new bottles (i.e., DO NOT USE), and discuss with DWM QC Analyst.

  - The preferred order for sample analysis is to start with the “oldest” samples first and proceed in the order in which they were collected.   QC samples (e.g., lab blank, lab duplicates, QC samples) and dilutions can be done at any time during the analysis.  If sample number is large and multiple samples are run at a time, be careful to put into its corresponding labeled tray.

  - NO DILUTION SAMPLES:  If no dilutions are to be performed on the sample, mix sample (as above) and immediately pour the sample into a new, sterile IDEXX vessel up to the 100 ml line.   If over the line, pour off to meet the 100 ml. meniscus, or use a sterile 1-10 ml pipette to pipette off to the 100 ml line.  (If IDEXX vessel is used to collect sample in the field, simply mix and then pour off/pipette off to the 100 ml. line).   Cap sample and proceed to 10.17.   See Table 3 below for dilution summary.                                                                                                 (NOTE:   Based on manufacturer specifications, the 100 ml line in the IDEXX vessel has a maximum measurement error of 2.5%.   Due to potential inconsistencies between analysts or systematic error, this approach may result in volume errors of greater than 2-3 mls.  Provided volumes errors are minimized (<2.5 %), this level of measurement accuracy and precision is currently deemed acceptable.)

  - 10 ML DILUTION SAMPLES: (e.g., for marine/brackish Colilert-18 samples; known or suspected “HOT” samples to be quantified at a level > 2420 org./100 ml.).    Pour 100 mls. of sterile DI water from a new WES-DI container into a new, sterile IDEXX vessel to the 100 ml. line (and pipette off as needed to 100 mls using a sterile 10 ml pipette.).   Now, using the same pipette, pipette off 10 mls to achieve 90 mls. of sterile water.  Then use the same pipette to pipette 10 mls. of  the original, mixed sample into the 90 mls. of sterile water.  Cap and mix as above.     See Table 3 below for dilution summary.                                                                                                                                                (NOTE:  An alternative method is to use pre-filled dilution bottles containing 90 mls. of phosphate buffer is under investigation).

  - 100 ML DILUTION SAMPLES:  (e.g., for marine/brackish Colilert-18 samples; for known or suspected “VERY HOT” samples to be quantified at a level >> 2420 org./100 ml.).  Pour 100 mls. of sterile DI water from a new WES-DI container into a new, sterile IDEXX vessel to the 100 ml. line (and pipette off as needed to 100 mls using a sterile 1 ml pipette.).   Now, using the same pipette, pipette off 1 ml. to achieve 99 mls. Then use the same pipette to pipette 1 ml. of the original, mixed sample into the 99 mls. of sterile water.  Cap and mix as above.     See Table 3 below for dilution summary.

  - Tap each reagent packet before opening to ensure complete dispensing.   Open Colilert reagent packet and add entire contents to the 100 ml. sample (or diluted sample).  Cap and invert the bottle > 25 times until all the reagent is dissolved.  Let any foam settle.  Some stubborn particles may remain; these may dissolve during incubation.

  - After complete mixing as stated above, dispense sample into the Quanti-Tray.  Use one hand to hold the Quanti-Tray upright with the well side facing the palm.  Squeeze upper part of the Quanti-Tray so that it bends towards the palm.  Open the Quanti-Tray by pulling the foil tab away from the well side. Avoid touching the inside of the foil or tray.   (NOTE:  If Quanti-Tray is believed to have been contaminated at this point, discard the contaminated Quanti-tray and set up a new one, if possible, and note in notebook).  Pour the Colilert-treated sample directly into the Quanti-Tray while avoiding contact with the foil tab.  Tap the small wells to release any air bubbles.

  - Place the sample-filled Quanti-Tray onto the rubber tray carrier of the Quanti-Tray Sealer with well side  (plastic) of the Quanti-Tray facing down to fit into the carrier.

  - Seal the Quanti-Tray by moving the rubber tray into the sealer. The sealer will automatically take the rubber tray to dispense the sample into the wells and seal the Quanti-Tray.  Remove the Quanti-Tray, assuring that it is fully sealed.  (NOTE:  If the Sealer stops in the middle of processing a Quanti-Tray sample, use the black button (with the arrows) on the front of the Sealer to reverse the direction of the Quanti-Tray in the Sealer.  Allow the green light on the Sealer to go on before attempting to seal the sample again.)  If the sealer does not easily “accept” the tray, lift up the end of the tray slightly while continuing to push it into the sealer, so that the leading edge clears the opening of the sealer.

  - Reagent Lot QC samples (DWM QA Analyst only).   Run the following E. coli QC samples for each new lot of reagents when reagents are first received:  1) Negative Control. Prepare P. aeroginosa (or other) samples as for E. coli below.   Cap/tray/seal/incubate.   2) Positive Control.  Prepare a positive QC sample by pouring 100 mls into an IDEXX vessel (as above), then adding an E. coli pellet according to manufacturer’s directions, then adding Colilert reagent.   Cap/tray/seal/incubate.  Confirm correct analytical end-points.

  - Batch QC samples (Lab analysts).  For each lab batch of samples, Prepare LAB BLANK as follows (minimum one per batch):  Pour 100 mls. of sterile, deionized water into an IDEXX vessel, add reagent, cap and mix, pour into tray, seal and incubate.  Prepare LAB DUPLICATE as follows (minimum one per batch):  Using 250 ml. sample:  Mix sample as above and pour two 100 mls. aliquots into two separate IDEXX vessels, add reagents, cap and mix, pour into trays, seal and incubate.   Using a 120 ml. sample:  Mix sample (each time) and pipette two 10 mls. aliquots into two separate IDEXX vessels containing 90 mls. of sterile, deionized water, add reagents, cap and mix, pour into trays, seal and incubate.   Multiply results by 10.   Provided 100 mls. is left in the sample container, run “raw”, 100 ml (undiluted) sample also.  NOTE:  Previous to 2011, lab analysts also ran positive and/or negative QC for each lab batch, but this was discontinued in 2011.

  - Incubate all prepared Colilert Quanti-Trays in the incubator at 35 ±0.5 oC for 24 - 28 hours.  If Colilert-18 is used, incubate for 18-22 hours at 35 ±0.5 oC.   Record the dates and times in the lab notebook.  Set clock.   Avoid any opening and closing the incubator during use in order to maintain stable proper temperature.

  - After sample trays have been placed in the incubator, thoroughly clean work area.   First, place all “bio-waste” (items that have been in contact with sample water, including pipettes, sample bottles, dilution bottles, etc.) in pre-labeled, plastic bio-waste bags.  Keep the plastic sample and reagent bottles separate from other waste (separate bags).   Tie the bag(s) when only half-full and temporarily store in the designated bio-waste storage area in the lab.   Recycle or dispose (as appropriate) other materials that have not touched the samples.   Once the work area is clear, disinfect the work area.

  - After 24 hours (18 hours for Colilert-18), remove the Colilert trays from the incubator.  Record the time in the lab notebook. Set clock.   For each tray, count the number of wells that turn yellow (at least as strong as the yellow color of the comparator).  A yellow color that is equal to or deeper than that of the color comparator verifies that the sample is positive for total coliforms.  Note in the lab notebook the number of positive (yellow) wells in the Quanti-Tray. Use the MPN table (Appendix A) to then record the confirmed total coliform value as MPN/100 ml sample in the lab notebook.

  - To determine the E. coli result, expose each sample to 365 nm UV light by placing them one-by-one into the UV light viewing cabinet (or in a dark room wearing UV-protective glasses with a hand-held UV lamp held within 5 inches of the sample.  If the sample fluoresces (compare to comparator), then the sample is positive for E. coli.  Count the number of fluorescent wells and record in the lab notebook. Use the MPN table (Appendix a) to record the confirmed E-coli value as MPN/100 mL sample in the lab notebook.  NOTE: Empty wells do not affect the test interpretation as long as the entire sample is in the tray. The effect on the Most Probable Number (MPN) table is statistically insignificant. An empty or partially filled well is interpreted the same way as a full well.

  - For added QC or if the sample results are questionable after 24 hours of incubation (18 hours for Coliert-18), the sample may be incubated for an additional 4 hours (total of 28 hours for Colilert; 22 for Colilert-18), and rechecked for color and fluorescent reactions.  If the color intensifies, the sample is total coliform positive; if it does not, the sample is negative.  If an inoculated test is inadvertently incubated over 28 hours (22 for Colilert-18), the following guidelines apply:  1) lack of yellow color is a valid negative test, and 2) a yellow color after 28 hours is not valid and must be repeated.

  - Colilert and Colilert-18 results are definitive at 24-28 and 18-22 hours, respectively.  Any positives for both total coliform and E.coli observed before the minimum time and negatives observed after the maximum time are also valid.  Positive results after the maximum time are not valid.

  - After all MPN results have been recorded, review the lab notebook to ensure that all required data and metadata have been recorded completely and accurately.

  - Place used Colilert trays and other waste in “bio-waste” bags, tie securely when only ½ full and temporarily store in the lab in the designated, secured location.

  - Periodically, coordinate with WES for the delivery of the “bio-waste” to WES for autoclave sterilization and disposal at WES.

Enterolert:  (in addition to or changes from that above for Colilert)

  - Incubate all prepared Enterolert Quanti-Trays in the incubator at 41 ±0.5 oC for 24 - 28 hours.

  - Prepare negative and positive controls as indicated in Section 8.

  - Dilute marine (brackish to salt) water samples with sterile fresh water at least ten fold, per IDEXX Enterolert procedure.

  - There is no fluorescent end-point comparator for Enterolert.  Compare fluorescence in sample trays to negative control.

General Dilution Schemes (depending on sampling bottle used and without pre-filled dilution bottles):

Table 2:  Summary of General Dilution Procedures

| Sample bottle used | Flip-top locking HDPE (or other) | IDEXX vessel PS |
| --- | --- | --- |
|  | Procedures |  |
| 0X dilution MIX bottle | Mix, pour 100 mls into clear PS IDEXX vessel, pour off/pipette off to 100 mls if necessary, add reagent | Mix, pour off/pipette off for necessary dilutions and to 100 mls, then add reagent |
| 10X dilution bottle | Pour 100 mls. of sterile DI into IDEXX vessel and pour off/pipette off to 100 ml. line as necessary. Pipette off 10 mls. Pipette 10 mls. of mixed sample into the 90 mls. in the IDEXX vessel. Add reagent | Pour 100 mls. of sterile DI into a new IDEXX vessel and pour off/pipette off to 100 ml. line as necessary. Pipette off 10 mls. Mix sample and pipette off 10 mls into 90 ml IDEXX vessel. Add reagent |
| 100 X dilution bottle | Same as above for 10X but using 1 ml. Add reagent | Same as above for 10X but using 1 ml. Add reagent |
| * These procedures assume that the IDEXX vessel 100 ml line can be used as an accurate (+/- 2.5%) measure. |  |  |

11.0	PREVENTATIVE MAINTENANCE

11.1	Sealer:  Trained personnel (only) must inspect, clean and maintain the sealer according to the manufacturer’s recommendations (Appendix C).

  - Rubber Inserts:  Rinse with clean water; autoclave or clean with isopropyl alcohol or bleach.

12.0	DATA INTERPRETATION, ANALYSIS AND VALIDATION

For each sample batch analysis, draft sample and quality control data are reviewed, interpreted and validated using the following definitions and criteria.

Data Report:   The lab analyst shall count the number of positive Quanti-Tray cells for each sample/tray, and use the MPN table to obtain the Most Probable Number (MPN) per 100 mls.  MPN results will be multiplied by the dilution factor as needed to obtain the final draft result.  The lab analyst shall then generate a brief draft results lab report for separate peer review.

Dilutions:  Multiply MPN value (using IDEXX MPN table) by the dilution factor used for that sample to get the sample result.

Peer Review and Preliminary Approval:   A separate lab report reviewer shall be designated to evaluate the accuracy and completeness of the lab report.    The reviewer shall discuss any problems, concerns and issues with the lab analyst as needed, and if all data appear OK, preliminarily approve the MPN results as DRAFT data.

Data Validation:  DWM’s QA/QC Analyst shall perform more rigid and comprehensive data validation for larger data sets (e.g., annual projects), including compiled E. coli MPN data (by project), in relation to project-specific data quality objectives (DQOs).

Table 3:  Colilert (and Colilert-18 Reactions)

| Reaction | Result or interpretation |
| --- | --- |
| Yellow | Total Coliform positive |
| No color or indeterminant | Negative |
| Fluorescent (yellow) | E. coli positive |
| Yellow color intensity at or near the comparator color | Inconclusive; re-incubate an additional 4 hours |
| > 28 (22) hours, positive | Not valid |
| > 28 (22) hours, negative | Valid |
| < 24 (18) hours, positive | Valid |
| < 24 (18) hours, negative | Not valid |

Table 4:  Enterolert Reactions

| Reaction | Result |
| --- | --- |
| Blue fluorescence in 24-28 hours | Enterococci positive |
| No fluorescence in 24-28 hours | Enterococci negative |
| Blue fluorescence < 24 hours | Enterococci positive |
| No fluorescence beyond 28 hours | Enterococci negative |

Table 5:  QC Interpretations (Colilert and Enterolert)

| Quality Control Organism | Yellow | Fluorescent |
| --- | --- | --- |
| Pseudomonas sp. (Non-fluorescent strain) | No | No |
| Klebsiella pneumonia, Enterobacter aerogenes or Enterobacter cloacae | Yes | No |
| E. coli | Yes | Yes (yellow) |
| Enterococcus faecium | --- | Yes (blue) |
| Serratia marcescens | --- | No |

Table 6:   QC Validation Criteria

| Quality Control | Frequency | Acceptance Criteria (control limits) | Corrective Action |
| --- | --- | --- | --- |
| Use of expired reagents and/or QC organisms | NA | Any and all >1 year passed exp date | Qualify data (m) Censor data (m) |
| Positive Culture Control (E. coli) | 5% | Positive | If not ALL positive, retest reagent lot. |
| Negative Control (sterile water) | 5% | Negative | If not ALL negative, retest reagent lot. |
| Sample storage | Every sample | Samples are analyzed within 8 hours of collection and are stored at  4C from time of collection to time of analysis | Qualify or censor data using H or P qualifiers (exceeded holding time or method of preservation not followed) |
| Lab duplicate (precision) | Min. one per batch per analyst and ~5% of samples | Range of logs = 0-TBD (based on mean range of logs + 3 SD for all lab duplicates , OR <30%RPD (<50 MPN); <20% (50-500 MPN); <10 % (500-5000 MPN); < 5% (>5000 MPN) for log10 transformed data (this DQO can also be applied to field duplicate evaluation) | Discard batch results outside control limits |
| Lab blank (bias) | Min. one per batch per analyst and ~5% of samples | < MDL | Qualify or censor data |

Red italics = proposed; 12/2008

13.0	METHOD PERFORMANCE

13.1	The method detection limit for Colilert, Colilert-18 and Enterolert has been determined to be a MPN of 1 colony-forming unit (CFU) per sample volume or dilution tested.  The DWM Lab minimum reporting limit (MRL) for data generated through use of this SOP has been chosen to be a MPN of 1 CFU per 100 mls.

13.2	Use of this method by DWM will be evaluated on a periodic basis to ensure consistent generation of quality-controlled data.

14.0	POLLUTION PREVENTION

14.1	In order to avoid or minimize waste due to non-use, the quantity of media and chemicals purchased should be based upon expected usage during its shelf life.

  - Actual preparation volumes should reflect anticipated usage and stability.

  - Recyclable materials, such as Type 6 polystyrene sterile sample bottles, should be used whenever possible.

  - Generation of wastewater will be minimized.  Discharge of any wastewater from sample preparation and analysis to the sanitary sewer will be minimized and does not pose a pollution potential.

  - All solid waste and “bio-waste” will be disposed of properly.

15.0	WASTE MANAGEMENT

  - All materials that have come in contact with samples (trays, bottles, pipettes, etc.) shall be bagged as “bio-waste” and temporarily stored in 32 gallon, plastic cans in a designated, secured location at DWM, prior to transport to the MADEP WES Lab in Lawrence, MA. for proper disposal (autoclave sterilization, plastics recycling and solid waste disposal).   See 15.3 also.

  - DWM’s  waste disposal procedures are consistent with State regulations (State Sanitary Code Title VIII, 105 CMR 480.00 and Solid Waste Regulations, 310 CMR 19.000) governing the disposal of “infectious waste”.  Infectious waste is considered a “special waste”, not a regulated hazardous waste.  It can be shipped to another location for treatment and disposal.

  - The used Colilert trays and other non-recyclable trash to be autoclaved shall be placed in separate, labeled “bio-waste” bags at DWM for transport to WES for sterilization and solid waste disposal.  The sample bottles and reagent bottles (recycled at WES) shall be bagged as “bio-waste” separately from the trays.   This will enable WES to autoclave and dispose of all materials properly.

  - DWM waste management procedures for Colilert shall adhere to all applicable aspects of DWM’s Lab Safety Plan.  This includes use of proper gloves, use of waste receipt forms when delivering waste to WES, and provisions for lab audits/inspections to ensure compliance and remedial training as necessary.

16.0	REFERENCES

16.1	Standard Methods for the Examination of Water and Wastewater, 20th Edition, 1998. (SM 9323B)  American Public Health Association, American Water Works Association, and Water Environment Federation, Washington, DC

  - http://www.state.ma.us/dep/bwp/dhm/files/infwaste.htm

  - http://www.idexx.com/

  - WES SOP for SM 9223:  Enzyme Substrate Coliform Test, Presence-Absence Procedure for Analysis of Potable Water Samples, Revision 2.0, 2/2003.

  - Total Coliform and E. coli by the Enzyme Substrate SOP, Ma. Water Resource Authority, Southboro Lab, 5/2003

  - Standard Operating Procedure for Bacteria Surveys, Sample Collection and Analysis in the ORSANCO Mobile Water Quality Laboratory, undated

17.0	APPENDICES

# Appendix A:   	MPN Tables for Colilert and Enterolert (by reference)

# Appendix B:	MSDS Sheets   (by reference; in DWM Labs and at http://www.idexx.com/ )

# Appendix C:	Sealer User’s Manual and Maintenance Instructions (by reference)

Appendix D:	Quanti-Cult Procedure  (by reference)

Appendix E:	Lab Notebook Example

Appendix F:	Incubator Temperature Log

Appendix G:	Waste Receipt Form

# Appendix E:   	Lab Notebook Example

| Analyst | Project | Test | Lab ID | OWM ID | Sample Date | Sample Time | Sample Receipt Date & Time | Analysis Date | Analysis Start Time | Media Lot # | Dilution | Incubator # | Incubation Start Time & Temp | Incubation End Time & Temp | UV Reader | # of large yellow wells | # of small yellow wells | Total Colifrm MPN | # of large fluores wells | # of small fluores wells | E. coli MPN |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| J. Bush | Blackstone Source Tracking | Colilert | L04-01-01 | 51-587 | 2/12/04 | 12:45 | 2/12; 1:50 | 2/12/04 | 2:24 | 368-01 | 0 | 1 | 4:35; 35.5 | 4:45; 35.0 | R. Chase | 14 | 14 | 67 | 7 | 7 | 34 |
|  |  |  | L04-01-02 | 51-587 | 2/12/04 | 12:45 |  | 2/12/04 | 2:26 | 368-01 | 10 |  | 4:35; 35.5 | 4:45; 35.0 |  | 2 | 2 | 8 | 1 | 1 | 5 |
|  |  |  | L04-01-03 | 51-588 | 2/12/04 | 12:55 |  | 2/12/04 | 2:29 | 368-01 | 0 |  | 4:35; 35.5 | 4:45; 35.0 |  | 37 | 37 | 287 | 26 | 26 | 211 |
|  |  |  | L04-01-04 | 51-588 | 2/12/04 | 12:55 |  | 2/12/04 | 2:31 | 368-01 | 10 |  | 4:35; 35.5 | 4:45; 35.0 |  | 4 | 4 | 27 | 4 | 4 | 27 |
|  |  |  | L04-01-05 | 51-588 | 2/12/04 | 12:55 |  | 2/12/04 | 2:35 | 368-01 | 100 |  | 4:35; 35.5 | 4:45; 35.0 |  | 0 | 0 | 0 | 0 | 0 | 0 |
|  |  |  | L04-01-06 | 51-589 | 2/12/04 | 1:25 |  | 2/12/04 | 2:45 | 368-01 | 0 |  | 4:35; 35.5 | 4:45; 35.0 |  | 9 | 9 | 49 | 5 | 5 | 32 |
|  |  |  | L04-01-07 | 51-590 | 2/12/04 | 1:25 |  | 2/12/04 | 2:47 | 368-01 | 10 |  | 4:35; 35.5 | 4:45; 35.0 |  | 2 | 2 | 8 | 2 | 2 | 8 |
|  |  |  | L04-01-07D | 51-590 Lab Dup | 2/12/04 | 1:25 |  | 2/12/04 | 2:48 | 368-01 | 10 |  | 4:35; 35.5 | 4:45; 35.0 |  | 3 | 3 | 15 | 2 | 2 | 8 |
|  |  |  | L04-01-08 | Lab Blank | 2/12/04 | 2:58 |  | 2/12/04 | 3:05 | 368-01 | 0 |  | 4:35; 35.5 | 4:45; 35.0 |  | 0 | 0 | 0 | 0 | 0 | 0 |
|  |  |  | L04-01-09 | E. coli + | 2/12/04 | 2:59 |  | 2/12/04 | 3:10 | 368-01 | 0 |  | 4:35; 35.5 | 4:45; 35.0 |  | 59 | 59 | --- | 59 | 59 | --- |

# Appendix F:   	Incubator Temperature Log

TEMPERATURE LOG

| Date | Time | Incubator TEMP, deg. C | TEMP (other) | Temp Dial Setting | Initials | Comments |
| --- | --- | --- | --- | --- | --- | --- |
| 4/5/04 | 9:15 | 35.5 | 35.0 | 1.6 | rc |  |

* Record before, during and after incubations.    See Jeff Smith, Richard Chase or Joan Beskenis if adjustment needed.

# Appendix G:   	Waste Receipt Form for “Bio-Waste”

WASTE RECEIPT FORM

GENERATOR  _______________________   	      ID # M _____________________________

RECEIVING FACILITY ___________________________          ID # M ______________________

NAME OF WASTE DELIVERED: _________________________            AMOUNT: ____________

SIGNATURE OF PERSON DELIVERING WASTE: ____________________          DATE: _______

SIGNATURE OF PERSON RECEIVING WASTE: _____________________          DATE: _______

GENERATOR COPY

WASTE RECEIPT FORM

GENERATOR  _______________________   	      ID # M _____________________________

RECEIVING FACILITY ___________________________          ID # M ______________________

NAME OF WASTE DELIVERED: _________________________            AMOUNT: ____________

SIGNATURE OF PERSON DELIVERING WASTE: ____________________          DATE: _______

SIGNATURE OF PERSON RECEIVING WASTE: _____________________          DATE: _______

Receiving Unit Copy

# Appendix H:   	Colilert® QuickGuide (these procedures also generally apply to Enterolert)

- Two hours prior to anticipated delivery of samples, turn on incubator and record temperature to verify 35.0 deg. C (=/- 0.5 C).

- Prepare all work areas by cleaning with lab disinfectant.

- When samples arrive, transfer custody of samples from crew to analyst using sign-off procedures (COC form)

- Allow samples to warm up to near room temperature.

- Turn on PC and open E.coli EDD template.  Fill out electronic lab notebook with sample information and per instructions in the “read me” tab.  This will generate the paper lab notebook page.

- Prepare counter work area with the necessary dilution bottles, reagents, trays, pipettes, etc., leaving sufficient space to prepare samples for incubation.  With a waterproof wax pencil (so as not to tear the paper on the back of the Quanti-Tray) gently mark the Quanti-Trays with the lab identification number of each sample. Inspect each IDEXX vessel for inside cleanliness; if particles or other contamination is observed, place the container aside with other contaminated, new bottles (i.e., DO NOT USE), and discuss with DWM QC Analyst.

- Turn on Quanti-Tray sealer to allow preheating (usually takes around 10 minutes).  Sealer is ready to use when green light is on.

- For each sample (including dilutions), invert sample bottle 20-25 times to ensure complete mixing of the sample prior to taking aliquots (for dilutions).  DO NOT TOUCH THE INSIDES OF CONTAINERS OR SAMPLE WATER.  The preferred order for sample analysis is to start with the “oldest” samples first and proceed in the order in which they were collected.   QC samples (e.g., lab blank, lab duplicates, QC samples) and dilutions can be done at any time during the analysis.

- NO DILUTION SAMPLES:  If no dilutions are to be performed on the sample, mix sample (as above) and immediately pour off sample water to the 100 ml line (bottom of meniscus).   Tap each reagent packet before opening to ensure complete dispensing.   Open Colilert reagent packet and add entire contents to the 100 ml. sample (or diluted sample).  Mix thoroughly until no more reagent particles.  Pour into pre-labeled tray.   Seal in sealer.

- 10 ML DILUTION SAMPLES:  Pour 100 mls. of sterile DI water (sterilized WES-DI water) into a new, sterile IDEXX vessel to the 100 ml. line (and pipette off as needed to 100 mls using a sterile 10 ml pipette.).   Now, using the same pipette, pipette off 10 mls to achieve 90 mls. of sterile water.  Then use the same pipette to pipette 10 mls. of the original, mixed sample into the 90 mls. of sterile water.  Cap and mix.     Add reagent.  Mix thoroughly until no more reagent particles.  Pour into pre-labeled tray.   Seal in sealer.

- 100 ML DILUTION SAMPLES:  Pour 100 mls. of sterile DI water into a new, sterile IDEXX vessel to the 100 ml. line (and pipette off as needed to 100 mls using a sterile 1 ml pipette.).   Now, using the same pipette, pipette off 1 ml. to achieve 99 mls. Then use the same pipette to pipette 1 ml. of the original, mixed sample into the 99 mls. of sterile water.  Cap and mix.   Add reagent.  Mix thoroughly until no more reagent particles.  Pour into pre-labeled tray.   Seal in sealer.

- Use of Trays.   After complete mixing as stated above, dispense sample into the Quanti-Tray.  Use one hand to hold the Quanti-Tray upright with the well side facing the palm.  Squeeze upper part of the Quanti-Tray so that it bends towards the palm.  Open the Quanti-Tray by pulling the foil tab away from the well side. Avoid touching the inside of the foil or tray.  Tap the small wells to release any air bubbles.

- Use of Sealer.  Place the sample-filled Quanti-Tray onto the rubber tray carrier of the Quanti-Tray Sealer with well side  (plastic) of the Quanti-Tray facing down to fit into the carrier.  Seal the Quanti-Tray by moving the rubber tray into the sealer. The sealer will automatically take the rubber tray to dispense the sample into the wells and seal the Quanti-Tray.  Remove the Quanti-Tray, assuring that it is fully sealed.  (NOTE:  If the Sealer stops in the middle of processing a Quanti-Tray sample, use the black button (with the arrows) on the front of the Sealer to reverse the direction of the Quanti-Tray in the Sealer.  Allow the green light on the Sealer to go on before attempting to seal the sample again.)  If the sealer does not easily “accept” the tray, lift up the end of the tray slightly while continuing to push it into the sealer, so that the leading edge clears the opening of the sealer.

- Reagent Lot QC samples (DWM QA Analyst only).   Run the following E. coli QC samples for each new lot of reagents:  1) Negative Control. Prepare P. aeroginosa (or other) samples as for E. coli below.   Cap/tray/seal/incubate.   2) Positive Control.  Prepare a positive QC sample by pouring 100 mls into an IDEXX vessel (as above), then adding an E. coli pellet according to manufacturer’s directions, then adding Colilert reagent.   Cap/tray/seal/incubate.  Confirm correct analytical end-points.

- Batch QC samples (Lab analysts).  For each lab batch of samples, Prepare LAB BLANK as follows (minimum one per batch):  Pour 100 mls. of sterile, deionized water into an IDEXX vessel, add reagent, cap and mix, pour into tray, seal and incubate.  Prepare LAB DUPLICATE as follows (minimum one per batch):  Using 250 ml. sample:  Mix sample as above and pour two 100 mls. aliquots into two separate IDEXX vessels, add reagents, cap and mix, pour into trays, seal and incubate.   Using a 120 ml. sample:  Mix sample (each time) and pipette two 10 mls. aliquots into two separate IDEXX vessels containing 90 mls. of sterile, deionized water, add reagents, cap and mix, pour into trays, seal and incubate.   Multiply results by 10.   Provided 100 mls. is left in the sample container, run “raw”, 100 ml (undiluted) sample also.  NOTE:  Previous to 2011, lab analysts also ran positive and/or negative QC for each lab batch, but this was discontinued in 2011.

- Incubate all prepared Colilert Quanti-Trays in the incubator at 35 ±0.5 oC for 24 - 28 hours.  If Colilert-18 is used, incubate for 18-22 hours at 35 ±0.5 oC.   Record the dates and times in the lab notebook.

- After sample trays have been placed in the incubator, thoroughly clean work area.   First, place all “bio-waste” (items that have been in contact with sample water, including pipettes, sample bottles, dilution bottles, etc.) in pre-labeled, plastic bio-waste bags.  Keep the plastic sample and reagent bottles separate from other waste (separate bags).   Tie the bag(s) when only half-full and temporarily store in the designated bio-waste storage area in the lab.   Recycle or dispose (as appropriate) other materials that have not touched the samples.   Once the work area is clear, disinfect the work area.

- After 24 hours (18 hours for Colilert-18), remove the Colilert trays from the incubator (and shut off).  Record the time in the lab notebook.  For each tray, count the number of wells that turn yellow (at least as strong as the yellow color of the comparator).  A yellow color that is equal to or deeper than that of the color comparator verifies that the sample is positive for total coliforms.  Note in the lab notebook the number of positive (yellow) wells in the Quanti-Tray.

- To determine the E. coli result, expose each sample to 365 nm UV light by placing them one-by-one into the UV light viewing cabinet (or in a dark room wearing UV-protective glasses with a hand-held UV lamp held within 5 inches of the sample.  If the sample fluoresces (compare to comparator), then the sample is positive for E. coli.  Count the number of fluorescent wells and record in the lab notebook.

- Colilert and Colilert-18 results are definitive at 24-28 and 18-22 hours, respectively.  Any positives for both total coliform and E.coli observed before the minimum time and negatives observed after the maximum time are also valid.  Positive results after the maximum time are not valid.

- After all MPN results have been recorded and entered into the electronic EDD, review the lab notebook to ensure that all required data and metadata have been recorded completely and accurately.  Save (as lab batch ID) and print EDD sheet.  Give printout to QC Analyst. Keep working printout with manually-recorded results in paper lab notebook.

- Place used Colilert trays and other waste in “bio-waste” bags, tie securely when only ½ full and temporarily store in the lab in the designated, secured location.

Colilert (and Colilert-18 Reactions)

| Reaction | Result or interpretation |
| --- | --- |
| Yellow | Total Coliform positive |
| No color or indeterminant | Negative |
| Fluorescent (yellow) | E. coli positive |
| Yellow color intensity at or near the comparator color | Inconclusive; re-incubate an additional 4 hours |
| > 28 (22) hours, positive | Not valid |
| > 28 (22) hours, negative | Valid |
| < 24 (18) hours, positive | Valid |
| < 24 (18) hours, negative | Not valid |

QC Interpretations (Colilert and Enterolert)

| Quality Control Organism | Yellow | Fluorescent |
| --- | --- | --- |
| Pseudomonas sp. (Non-fluorescent strain) | No | No |
| Klebsiella pneumonia, Enterobacter aerogenes or Enterobacter cloacae | Yes | No |
| E. coli | Yes | Yes (yellow) |
| Enterococcus faecium | --- | Yes (blue) |
| Serratia marcescens | --- | No |