Massachusetts Department of Environmental Protection Division of Watershed Management

DRAFT

STANDARD OPERATING PROCEDURE

Screening-Level Microcystins

CN 230.0

November, 2012

| Prepared by: |  | Date: |  |
| --- | --- | --- | --- |
|  | ###### Joan Beskenis, Environmental Analyst |  |  |
| Approved by: |  | Date: |  |
|  | ###### Richard Chase, Quality Assurance Analyst |  |  |
| Approved by: |  | Date: |  |
|  | ###### Arthur Johnson, Monitoring Coordinator |  |  |
| Approved by: |  | Date: |  |
|  | ###### Dennis Dunn, Program Supervisor |  |  |

### * see pdf version for valid signatures

- SCOPE AND APPLICATION

2.0       SUMMARY

3.0       SAFETY CONSIDERATIONS

- Solution is 1.0 N Hydrochloric acid. Handle carefully.

- Wear disposable gloves and safety glasses

4.0       SAMPLE COLLECTION, PRESERVATION AND HANDLING

The type of water sample collection method depends, in part, on project protocols.  Sample types can range from manual grabs to depth-integrated or fixed-depth samples.  Use clean glass containers for all samples.

Grab samples are obtained from a boat or by wading in-stream where possible. Boat samples should be collected from the upstream/windward side of the boat.  Manual grab samples are collected by holding the base of the container and plunging the capped container beneath the surface.  The container neck should point slightly upward and the mouth should be directed toward the current.  If there is no current, an artificial one can be created by pushing a container forward horizontally in a direction away from the hand.  Open the sample container 15-30 cm below the surface to avoid collecting surface material.  Fill the container to the shoulder to allow enough air space for mixing.

For lake sampling, a pre-cleaned, labeled 250 ml HDPE container is used in the collection of an integrated sample of the water column.   An integrated sample is collected using thin walled polyethylene tubing (3/8 inch internal diameter/0.9 cm).  The tubing is lowered down to the desired depth (3 times the secchi disk reading), the top is crimped at the surface and then it is pulled out of the water.  The tube end is held near the mouth of the sample container, but held so that the tubing does not contact the inside of the container. De-crimping allows the water to drain out; the sampling procedure is repeated until at least 300 mL are collected.  Alternatively, a Van Dorn bottle can be used to collect samples at fixed-depths.

The sample containers should be stored on ice in a cooler until delivery at CERO-Worcester.

5.0      APPARATUS, EQUIPMENT, MATERIALS

- DIW (Barnstead system, DWM-CERO)

- Timer

- Pipettes

- Test tube rack

- Envirologix Qualitube Kit for Microcystin.

  - Kit Detection Range:  0.5-3 ppb.

  - Method Limit of Detection:  0.3 ppb

#### 6.0       REAGENTS

- Envirologix Qualitube Kit for Microcystin.  Store in 4-8 deg. C fridge

- De-ionized water (DIW) is prepared at DWM using a Nanopure Ion Exchange Filtration System.

7.0      CALIBRATION

N.A.

- PROCEDURE

- Allow all reagents to reach room temperature before beginning (at least 30 minutes with un-boxed tubes and reagents at room temperature – do not remove tubes from bag with desiccant until they have warmed up).

- Organize all samples and reagents so that steps 1 and 2 can be performed in 3 minutes or less.  Do not run more than 6 tubes at a time.

- Rapidly add 5 drops of Microcystin Assay Diluent to each tube in the assay.

- Using the sample pipette provided, immediately add two drops of 0.5 ppb Microcystin Calibrator to the first tube. Add two drops of 3.0 ppb Microcystin Calibrator to the second tube. Add two drops of sample to each of the subsequent tubes, up to a total of 4 samples. Do not add Microcystin-enzyme Conjugate in this step.

- Thoroughly mix the contents of the tubes by moving the tube holder in a rapid circular motion on flat surface for a full 20-30 seconds.

- Incubate tubes at ambient temperature for 5 minutes.

- Add 5 drops of Microcystin-enzyme Conjugate to each tube. Do not empty the tube contents or wash the tubes at this time. Thoroughly mix the contents of the tubes as in step 3.

- Incubate tubes at ambient temperature for 20 minutes.

- After incubation, vigorously shake the contents of the tubes into a sink or other suitable container. Flood the tubes completely with cool tap water, then shake to empty. Repeat this wash step three times. Invert the tubes on a paper towel and tap to remove as much water as possible.

- Add 10 drops of Substrate to each tube. Thoroughly mix the contents of the tubes, as in step 3. Incubate substrate in tubes for 10 minutes at ambient temperature.

- NOTE: If blue color does not develop in the 0.5 ppb Calibrator tube, the assay is invalid and should be repeated.

- Caution: Stop Solution is 1.0 N Hydrochloric acid. Handle carefully.

- This assay is designed to be read visually with un-stopped tubes (blue solution). If tubes are to be read using a tube photometer, pipette 0.7 mL of Stop Solution into each tube and mix thoroughly. This will turn the tube contents yellow.

- NOTE: Read the tubes within 30 minutes of the addition of Stop Solution.

- Interpret the results of un-stopped tubes immediately following the 10 minute substrate incubation.

How to Interpret the Results Reading Tubes Visually:

1. Compare the intensity of the blue color of each sample tube to the intensity of the blue color in the 0.5 and 3.0 ppb calibrator tubes.

2. Score each sample tube as having less than, more than or equal color to the two calibrator tubes.

3. Use Table to determine the level of microcystin in the samples.

| Samples with blue color … | Contain … |
| --- | --- |
| Darker than the blue color of 0.5 ppb Calibrator | Less than 0.5 ppb Microcystins |
| Between the blue color of 0.5 ppb and 3.0 ppb Calibrator | Between 0.5 and 3.0 ppb Microcystins |
| Lighter than the blue color of 3.0 ppb Calibrator | More than 3 ppb Microcystins |

- QUALITY CONTROL

A laboratory duplicate is run on every tenth sample.  The precision for these samples will be determined by calculating their relative percent difference (RPD).

RPD= (sample- duplicate)/(average of sample and duplicate)*100.  The RPD should not exceed 20%.

A laboratory blank………….

- INTERFERENCES

Possibly humic acids >100 ppm.

- PREVENTIVE MAINTENANCE

- CORRECTIVE ACTIONS

13.0        WASTE AND POLLUTION PREVENTION

14.0       REFERENCES

www.envirologix.com

15.0      APPENDICES

Appendix A:  Envirologix Qualitube Kit Instructions

Appendix A:  Envirologix Qualitube Kit Instructions