Massachusetts Department of Environmental Protection Division of Watershed Management

STANDARD OPERATING PROCEDURE

Hanna ammonia high range portable photometer HI 96733

CN 399.0

April, 2015

| Prepared by: |  | Date: |  |
| --- | --- | --- | --- |
|  | ###### Jennifer Sheppard, Environmental Analyst |  |  |
| Approved by: |  | Date: |  |
|  | ###### Richard Chase, Quality Assurance Analyst |  |  |
| Approved by: |  | Date: |  |
|  | ###### Arthur Johnson, Monitoring Coordinator |  |  |
| Approved by: |  | Date: |  |
|  | ###### Kim Groff, Program Supervisor |  |  |

# TABLE OF CONTENTS

Scope and Application

Components

Specifications

Storage and Transport of Unit

Calibration

Quality Assurance, Quality Control

Data Validation and Management

Maintenance and Repair

References

Appendices

## SCOPE and Application

This SOP details how to use, maintain and care for the HANNA instruments ammonia high range portable photometer HI 96733 (hereafter known as the HI 96733) as its use relates to MassDEP-SEROBST water monitoring activities.

The HI 96733 is an auto diagnostic portable microprocessor meter; it measures the ammonia nitrogen (NH3-N) content in water samples in the 0.0 to 50.0 mg/L (ppm) range. It has an advanced optical system based on a special tungsten lamp and a narrow band interference filter. Cuvettes made of optical glass are placed into the measurement cell. The instrument is factory calibrated and the electronic and optical design minimizes the need of frequent calibration. The instrument is splash waterproof and the tungsten lamp and filter units are protected from dust or dirt by a transparent cup. This makes the instruments fulfill field applications. The “Nessler” method is used, whereby the reaction between ammonia and the reagents causes a yellow tint in the sample. The reagents are liquid form and are supplied in bottles. The amount of reagent is precisely dosed to ensure the maximum repeatability.

Measurement of ammonia is a useful component of the Bacteria Source Tracking (BST) program.  In combination with a measurement of potassium and expressed as a ratio “ammonia/potassium”, it can be an indicator of pollutant sources and help to distinguish between types of illicit discharges (Pitt, 2004).

At this time, ammonia meter data collected in the field by the BST program is not currently managed by DWM database staff.

## COMPONENTS

The necessary accessory equipment to operate and maintain the HI 96733 Ion selective meter includes:

- two sample cuvettes and caps

- carrying case

- 1ml syringe

- disposable plastic pipettes

- HI 93733B-0 Ammonia Reagent B

- HI 93733A-0 Nessler Reagent

- The meter takes one 9 volt battery.

# SPECIFICATIONS

Meter resolution and accuracy, as provided by the manufacturers, are provided below. These specifications represent a baseline of expected performance and for comparison to results.

| Mode | NH3-N |
| --- | --- |
| Range | 0.0 to 50.0 mg/L |
| Resolution | - mg/L for measurement 0.1 mg/L for calibration and validation |
| Accuracy | ±0.5 mg/L ±5% of reading @ 25oC |

#### Storage of MEter

When not in use, the HI 96733 shall be kept in its carry case in the SERO lab in an appropriate dedicated space. The carry case shall be closed at all times when the meter is not in use.

The meter, (especially the cuvette holder) shall be kept clean and free of dust, sand, grit, and water.

Calibration

The instrument is factory calibrated and the electronic and optical design minimizes the need of frequent calibration. If calibration must be performed, then such activities shall be carried out by SEROBST staff. All calibration data shall be stored in lab notebooks.

The equipment and supplies listed below are necessary for calibration of the HI 96733:

- HANNA CAL CHECK Standards

- CAL CHECK Standard Cuvettes

At this time SEROLAB does not stock the supplies necessary for calibration of the HI 96733. They are extremely toxic substances, which require the use of a hood. SEROLAB is not currently equipped to use such substances.

#### Material Safety and Waste Management

The Nessler reagent A is highly toxic, toxic by inhalation, in contact with skin and if swallowed. It is also very toxic to aquatic organisms. See MSDS (attachment 2) for detailed instructions on the handling of this substance and procedures to follow in case of spills. The MSDS states that this substance should not be allowed to enter waters, waste waters or soil, consequently the waste by-products of testing a sample for ammonia are disposed of into a plastic container with a screw top lid that is clearly marked and kept at the SERO Lab. At the end of the sample season this container is transported to WES (The State Lab) for safe disposal.

#### use of the HI 96733

This instrument is used in a laboratory setting at all times.

General tips for an accurate measurement of ammonia

To ensure best accuracy:

- Use active carbon and prior filtration to remove any color or suspended matter from sample to be measured, since it may cause interference.

- When filling the cuvette, the base of the meniscus must come to the same level as the 10ml mark.

- When using the dropper bottle of Nessler reagent:

  - Tap the dropper first on the table and wipe the outside of the dropper with a cloth.

  - Always keep the dropper bottle in a vertical position while dosing the reagent

- Make sure all debris is cleared from sample before measurement

- To avoid leaks and obtain more accurate measurements, use the plastic HPDE stopper to close the cuvette followed by the black cap.

- Each time the cuvette is used the cap must be tightened to the same degree

- Before measurements wipe cuvette outside thoroughly with a lint free cloth so that it is dry and free of fingerprints oil and dirt.

- Remove any bubbles in the sample before measurement by tapping or gently swirling.

- Do not let the reacted sample stand too long after reagent is added, or accuracy will be lost

- It is possible to take multiple readings in a row, but it is recommended to take a new zero reading for each sample and to use the same cuvette for zeroing and measurement

- After the reading, immediately discard the sample, otherwise the glass might become permanently stained.

- All the reaction times reported in this manual are referred to 20oC (68F). As a general rule of thumb, they should be doubled at 10oC (50F) and halved at 30oC (86F).

Measurement Procedure

This takes place in two phases to compensate the meter for sample turbidity or color.

First the meter is zeroed using the unreacted sample, then after the reagents are added the reacted sample is measured.

- Turn meter on

- Zero the meter: Fill a cuvette with 1ml of unreacted sample by means of the syringe. Using the plastic pipette, add 9ml of ammonia reagent B up to the 10ml mark. Place the cap and swirl the solution to mix. Place cuvette into the holder and press the ZERO/CFM button. Once the display shows 0-0 the meter is ready for measurement.

- Remove cuvette from the holder and add 4 drops of Nessler Reagent. Swirl to mix.

- Place cuvette into the holder and press and hold READ/TIMER button for three seconds. The display will show a countdown prior to measurement. A beep will sound at the end of the countdown period. The instrument will display concentration in mg/l of ammonium ion (NH4+).

- To convert the reading to mg/l of ammonia (NH3), multiply by the factor 0.944.

#### QUALITY ASSURANCE, QUALITY CONTROL

A duplicate set of readings can be taken at a rate of once per day by taking a second aliquot of sample water from a sample bottle and going back through the same measurement procedure (as above). The duplicate readings can provide information on overall precision or repeatability of the in-situ measurements.    Most of any variation observed will be assumed to typically be due to natural variation, but this assumption may not hold in all cases, which may lead to inferences regarding quality control.

# DATA VALIDATION and MANAGEMENT

Users shall be responsible for recording ammonia measurements taken by the HI 96733 in a lab notebook.  HI 96733 data will be reviewed, validated and assessed for usability by DWM staff, including Regional Monitoring Coordinators.  Based on manufacturer accuracy specifications, as long as duplicate data is within the range (+/-) of probe accuracy, for the parameter in question (see specifications table above) then it will be considered acceptable for the current uses (i.e. screening purposes). If the difference between duplicate data is not within the range (+/-) of probe accuracy the data will be censored.

At this time analytical data collected in the field by the DWM-BST programs (other than bacteria) are not currently managed by DWM. Data from the HI 96733 will be transferred to computer data files as soon as possible after lab work is complete.

### Maintenance and Repair

There are no maintenance details listed in the instruction booklet for this meter. Consultation with technical personnel at HANNA Instruments may be required when there is uncertainty about a particular repair problem.

References

Pitt, Robert & The Center for Watershed Protection.  October 2004. Illicit Discharge Detection and Elimination-A Guidance Manual for Program Development and Technical Assessments.

Attachment 1:  Manufacturer Instruction Manual

http://www.hannainst.com/manuals/man96733_22_12_2010.pdf

Attachment 2:  MSDS

http://site.jjstech.com/pdf/Hanna-Instruments/SDS_HI%2093733A-0_2012-05-21.pdf

http://hannainst.com/sds/SDS_HI%2093733B-0_2013-06-27.pdf

http://www.hannainst.com/sds/SDS_HI%2096733B_2012-04-30.pdf

(For Hanna CAL CHECK Standards, not currently held at SEROLAB