STANDARD OPERATING PROCEDURE

#### For

#### USEPA METHOD 353.1

#### Determination of Nitrate + Nitrite-Nitrogen by the

#### Automated-Colorimetric, Hydrazine Reduction Method

| ## SOP #: EPA 353.1 | REVISION #: 1.2 | June 2003 |
| --- | --- | --- |
|  | Page 1 of 23 |  |
| This and all other DEA/WES SOP and QA documents are available (read/print only) to all WES employees on the WES server (w:\dea-qap\SOPs & QA Docs). It should be noted that the controlled SOP & QA documents are only those viewed on-line on the WES server. If this is a printed copy, it is an uncontrolled version and may not be the latest version currently in use. |  |  |

Massachusetts Department of Environmental Protection Division of Environmental Analysis Senator William X. Wall Experiment Station 37 Shattuck Street, Lawrence, MA 01843

| Prepared by: |  |  |  |  |
| --- | --- | --- | --- | --- |
|  |  |  | Date: | June 4, 2003 |
|  | ###### Maria E. Ruiz, Chemist |  |  |  |
| Revised by: |  |  | Date: | June 4, 2003 |
| Approved by: | ###### Maria E. Ruiz, Chemist |  |  |  |
|  |  |  | Date: |  |
| Approved by: | ###### James Sullivan, Laboratory Supervisor |  |  |  |
|  |  |  | Date: |  |
|  | ###### Robert Serabian, Quality Assurance Officer |  |  |  |
| Approved by: |  |  | Date: |  |
|  | ###### Oscar Pancorbo, Director |  |  |  |

TABLE OF CONTENTS

Page

LIST OF REVISIONS	3

LIST OF TABLES	4

1.0	SCOPE AND APPLICATION	5

2.0	SUMMARY OF METHOD	5

3.0	DEFINITIONS	5

4.0	INTERFERENCES	6

5.0	SAFETY	6

6.0	EQUIPMENT AND SUPPLIES	7

7.0	REAGENTS AND STANDARDS	7

8.0	SAMPLE COLLECTION, PRESERVATION, AND STORAGE	9

9.0	QUALITY CONTROL	9

10.0	CALIBRATION AND STANDARDIZATION	10

11.0	PROCEDURE	12

12.0	DATA ANALYSIS AND CALCULATIONS	17

13.0	METHOD PERFORMANCE	17

14.0	POLLUTION PREVENTION	17

15.0	WASTE MANAGEMENT	17

16.0	REFERENCES	18

17.0	TABLES AND VALIDATION DATA	19

# LIST OF REVISIONS

| Rev. # | Date | Description of Revision | Page # |
| --- | --- | --- | --- |
| 0 | March 2001 | None |  |
| 1.0 | October 2001 | Table 1 – updated Table 2 – updated Table 3 – renamed Table 4 New Table 3 added | 18 18 19 19 |
| 1.1 | November 2002 | Section 3.1 renumbered to 3.2 Section 3.2 – 3.12 renumbered to 3.4 - 3.14 New Section 3.1 and 3.3 Section 6.4 – new printer Hewlett-Packard Deskjet 692C Section 8.2 – “at 4C” added Section 8.3 added Section 9.1 – “LRB” changed to “Continuing Calibration Blank”, “after each batch of samples” changed to “every 10 samples Section 9.2: “A continuing calibration standard” changed to “The Instrument Performance Check Solution” Section 13.3 “three to five times” changed to “two to three times” Table 1 – Table 4 renumbered to Table 2 – Table 5 Table 5 updated New Table 1 added | 5 5 – 6 5 8 9 9 9 9 17 20 – 23 23 20 |
| 1.2 | June 2003 | Section 14.3 rewritten Section 15.1 rewritten Table 2 through Table 5 values updated | 19 19 21 – 23 |

# LIST OF TABLES

Page

TABLE 1.	Quality Control Elements and Acceptance Limits for EPA Method 353.1 Determination Of Nitrate, Nitrite Nitrogen Colorimetric- Automated, Hydrazine Reduction	19

TABLE 2.	Accuracy of Nitrate Nitrite Nitrogen Analysis in Water Samples by EPA Method 353.1	20

TABLE 3.	Precision of Nitrate Nitrite Nitrogen Analysis in Water Samples by  EPA Method 353.1	21

TABLE 4.	Accuracy of Nitrate Nitrite Nitrogen Analysis in Reagent Water (LFB) by EPA Method 353.1	22

TABLE 5.	Method Detection Limit (MDL) for Nitrate Nitrite Nitrogen Analysis in Reagent Water by EPA Method 353.1	22

# 1.0	SCOPE AND APPLICATION

1.1	This method is applicable to surface water, domestic and industrial wastes.

1.2	The applicable range of this method is 0.02 to 2.00mg/L nitrate- nitrite nitrogen.  The range may be extended with sample dilution.

1.3	It uses the principle of air-segmented continuous-flow analysis (CFA) for fully automatic sample analysis of liquid samples using wet-chemistry methods..  Samples are mixed with reagents in a continuously flowing stream.  The individual sample segments are kept separate by means of air bubbles.

1.4	Operate in STEADY STATE mode- characteristic of segmented flow system- where the absorbance of the reaction stream is not changing with time.

# 2.0	SUMMARY OF METHOD

2.1	Nitrate is reduced to nitrite by an alkaline solution of hydrazine sulfate containing a copper catalyst and the nitrite (that originally present plus reduced nitrate) is determined by diazotizing with sulfanilamide and coupling with N-(naphtyl)- ethylenediamine dihydrochloride to form a highly colored azo dye which is measured colorimetrically.

# 3.0	DEFINITIONS

3.1	Calibration Blank (CB).  A volume of reagent water fortified with the same matrix as the calibration standards, but without the analytes.

3.2	Calibration Standard (CAL)- A solution prepared from the dilution of stock standard solutions or from the primary dilution standard solution. The CAL solutions are used to calibrate the instrument response with respect to analyte concentration.

3.3	Instrument Performance Check Solution (IPC) – A solution of one or more method analytes used to evaluate the performance of the instrument system with respect to a defined set of criteria.

3.4	Laboratory Duplicates - Two aliquot of the same sample taken in the laboratory and analyzed separately with identical procedures.

3.5	Laboratory Fortified Sample Matrix (LFM) - An aliquot of an environmental sample to which known quantities of the method analytes are added in the laboratory.

3.6	Laboratory Reagent Blank (LRB) - An aliquot of reagent water that is treated exactly as a sample.  The LRB is used to determine if method analytes or other interferences are present in the laboratory environment, reagents, or apparatus.

3.7	Method Detection Limit (MDL) - The minimum concentration of an analyte that can be identified, measured and reported with 99% confidence that the analyte is greater than zero.

3.8	Standard Reference Material (SRM) - A solution of method analyte of known concentrations, that is obtained from a source external to the laboratory and different from the source of calibration standards.

3.9	Laboratory Fortified Blank (LFB) - An aliquot of reagent water or other blank matrices to which known quantities of the method analytes are added in the laboratory.

3.10	Stock Standard Solution (SSS) - A concentrated solution containing one or more method analytes prepared in the laboratory using assayed reference materials or purchased from a reputable commercial source.

3.11	CAS -Chemical Abstract Service.

3.12	STORET - Storage and Retrieval of Water-Related Data.

3.13	Material Safety Data Sheet (MSDS) - Written information provided by vendors concerning a chemical's toxicity, health hazards, physical properties, fire, and reactivity data including storage, spill, and handling precautions.

3.14	Linear Calibration Range (LCR) - The concentration range over which the instrument response is linear.

# 4.0	INTERFERENCES

4.1	Sample color that absorbs in the photometric range used for analysis will interfere.

4.2	Chloride, sulfide, ferric ion and phosphate ion interfere.

4.3	The apparent NO3 and NO2 concentrations varied  10 percent with concentrations of sulfide ion up to 10 mg/L.

# 5.0	SAFETY

5.1	Standard laboratory protective clothing and eye covering is required.

5.2	CAUTION: Hydrazine sulfate stock solution is toxic if ingested.  Mark container with appropriate warning.

5.3	Acidification of samples should be done in a fume hood.

5.4	All personnel handling environmental samples known to contain or to have been in contact with human waste should be immunized against known disease causative agents.

5.5	It is the responsibility of the user of this method to comply with relevant disposal and waste regulations.

5.6	All work performed under clean hood.

# 6.0	EQUIPMENT AND SUPPLIES

6.1	Analytical Balance- Capable of accurately weighing to the nearest 0.0001 g.

6.2	AutoAnalyzer 3 BRAN+LUEBBE.  Continuous- Flow Analytical System.The main system components include:

6.2.1	Compact  Sampler- BRAN + LUEBBE . Is a space-saving random access sampler, with 100-cup capacity.

6.2.2	AA3 Digital Colorimeter- 2 channel-BRAN + LUEBBE - Provide the detector module to the AA3 system as well as the control electronic for all AA3 modules. 	Colorimeter equipped with 1.0mm diameter and 10mm path length flow cell (vol. 8 µL) and 520 nm filter.  The standard detector is a high sensitivity photometer.

6.2.3	AA3 Pump.  A high precision peristaltic pump that pumps samples, reagents, diluents and air* through a series of flexible pump tubes.  By a steady process of first squeezing and then relaxing pressure on the tubes, the fluids are drawn into the system and pushed forward through the components on the Chem Tray until they arrive in the flowcell of the Digital Colorimeter.  * The air bubbles that are required to divide the sample and reagent streams into equally spaced segments, generated by a valve module that provide space for 4 air lines (silicone tubes).

6. 2.4	AA3 Chemistry Module. Contains all the components required for the reaction, e.g. mixing coils, heating bath 37ºC.  The components for one method are mounted on a so-called Chem Trays and can be used to determine two parameters simultaneously.

6.3	PREMIO Computer, INC – Model: PREMIO MTP 6C-PENTIUM II Mid-Tower.

6.3.1	Operation system: Windows 95.

6.3.2	Control software: AACE  A Bran + Luebbe’s dedicated continuos flow software package that control the AA3 modules.

6.4	Printer- Hewlett Packard Deskjet 692C

6.5	Labware.  All washed with a detergent solution, rinsed with tap water.  Prior to use, H2SO4 washed and reagent water rinsed.

# 7.0	REAGENTS AND STANDARDS

7.1	Color developing reagent:  To approximately 500 mL of reagent water add 200 mL concentrated phosphoric acid (sp gr. 1.834) and 10g sulfanilamide (H2NC6H4SO2NH2).  Dissolve completely.  Add 0.8g N-(1-Naphtyl) ethylenediamine  Dihydrochloride.  Dilute the solution to 1 liter with D.I water.  Add 1.0 mL BRIJ-35.  Store in a dark bottle and keep refrigerated.  STABILITY: ONE MONTH.

7.2	Copper sulfate stock solution:  Dissolve 3.0 g of cupric sulfate (CuSO4  5H2O) in reagent water and dilute to one liter.

7.3	Stock hydrazine sulfate:  Dissolve 30.0 g of hydrazine sulfate (N2H4.H2SO4) in 800 mL of reagent water.  Dilute to one liter.  This solution is stable for six months if stored in a tightly stoppered amber bottle.

7.4	Working Reductor solution:  Dilute 30 ml of stock hydrazine sulfate in about 900 mL with reagent water.  Add 10 mL of stock cupric sulfate solution.  Dilute to one liter with reagent water.  Store in an amber bottle.  STABILITY: ONE MONTH.   When first prepared, test the working reductor solution for efficient reduction capacity.  Run both a nitrite and nitrate standard (2.0 mg/L).  If the apparent nitrate concentration is much lower than the nitrite concentration, adjust the working reductor solution by 1mL aliquot additions of the stock hydrazine solution until the two standards are equal.  If too much stock hydrazine solution is added, the nitrite peak will appear lower than the nitrate peak.  When the proper concentration of hydrazine sulfate has been attained, no further adjustment is necessary.

7.5	Sodium hydroxide stock solution, 50% w/w:  Weigh  500g NAOH pellets.  In a large beaker weigh 500g reagent water.  Slowly, over an hour or more, add the NAOH pellets, waiting for the solution to cool somewhat between additions.  Store in a plastic container and make up as needed.

7.6	Sodium Hydroxide working solution:  Weigh 14.4 g (approx. 10.6 ml) of sodium hydroxide stock solution, 50% w/w.  Dilute to 1000 mL with reagent water.  Add 2 mL (40 drops) of Brij-35 and mix well. Store in a plastic container, make up as needed.

7.7	Stock nitrate solution (1000 mg/L):  Dissolve 7.218 g of KNO3, oven dried at 100-105ºC for 24 hours, in DI.water and dilute to 1 liter.  Add 1 mL chloroform  (CHCL3) as a preservative.  Stable for 6 months.  Keep stock solution in the refrigerator when not in use.

7.8	Standard Nitrate solution (100 mg/L): Dilute 100 mL of stock nitrate solution (7.7) to 1 liter.

1 mL = 0.1 mg N.

NOTE: This standard is prepared combined with NH3-N.

7.9	Stock nitrite solution (100 mg/L NO2-N): Dissolve 0.4926g of anhydrous NANO2 (24 hours in desiccator) in reagent water and QS to 1 liter.  Preserve with 2 mL chloroform.

1 ml = 0.1 mg NO2-N

7.10	Preparation of  Calibration Standards:  Using the standard nitrate-ammonia/N solution (7.8), prepare the following calibration standards in  100 ml volumetric flasks.

Needed are the following concentrations (mg/L): 10, 2.0, 1.5, 1.0, 0.5, 0.1, 0.05

One nitrite standard (2.0 mg/L) is compared to a nitrate standard (2.0 mg/L) to verify the efficiency of the reduction in every run.

Make up 100 mL of each:

10.0 mg/L:	10 mL of 100 mg/L, QS to 100 mL with reagent water.

2.00 mg/L:	2 mL of 100 mg/L, QS to 100 mL with reagent water.

1.50 mg/L:	15 mL of 10 mg/L, QS to 100 mL with reagent water

1.00 mg/L:	10 mL of 10 mg/L, QS to 100 mL with reagent water

0.50 mg/L:	5 mL of 10 mg/L, QS to 100 mL with reagent water

0.10 mg/L:	10 mL of 1 mg/L, QS to 100 mL with reagent water

0.05 mg/L:	5 mL of 1 mg/L, QS to 100 mL with reagent water

Calibration Blank

0.02 mg/L:	0.2 mL of 10 mg/L, QS to 100 mL with reagent water (used as MDL standard).

0.40 mg/L:	 4 mL of 10 mg/L, QS to 100 mL with reagent water (used as LFB).

0.06 mg/L:	0.6 mL of 10 mg/L, QS to 100 mL with reagent water (used as RDL)

7.11	Reagent Water:  ASTM Type I grade water.

7.12	SYSTEM WASH SOLUTION.  Use reagent water containing 2 mL/L Brij-35, 30% solution.

7.13	SPECIAL WASH SOLUTION:  Use a 1 N HCL solution (about 83 ml/L conc. hydrochloric acid).

NOTE:	Sampler wash solution does not contain wetting agent.

7.14	Brij-35

# 8.0	SAMPLE COLLECTION, PRESERVATION, AND STORAGE

8.1	If analysis can be made within 24 hours, samples should be preserved by refrigeration at 4ºC.  When samples must be stored for more than 24 hours, they should be preserved with 2 ml of sulfuric acid  (H2SO4) per liter and refrigerated.

8.2	Maximum holding time is 28 days for preserved samples pH <2 at 4C.

8.3	Samples should be collected in plastic or glass bottles.

# 9.0	QUALITY CONTROL

9.1	A Continuing Calibration Blank is run after calibration, every ten samples and at the end of the run.

9.2	The Instrument Performance Check Solution (a midrange check standard) is run after calibration, every 10 samples and at the END of the run.

9.3	Two QCS are analyzed after calibration and at the end of the run. Reading must be within ± 10%.

9.4	Duplicates and LFM are run every 10 samples or less.

9.5	A LFB is analyzed with each batch of samples analyzed.

9.6	After calibration a nitrite and nitrate standard (2.0 mg/L) is run to check reductor solution or efficient reduction capacity (100% reduction of nitrate to nitrite).  The two peaks should be of equal height.

9.7	A MDL standard is run with each run.

9.8	A RDL standard is run with each run.

# 10.0	CALIBRATION AND STANDARDIZATION

10.1	Remove reagents, 100 mg/L NO3-N and NH3-N combined stock standard, QCS and samples from refrigerator. Prepare reagents as needed.  Prepare working standards.

10.2	Rinse diluent-wash solution flask with reagent water and refill with DI and BRIJ –35* (30% solution) from BRAN + LUEBBE, part # T21-0110-04  (15 drops of BRIJ per 500 ml of reagent water). Place tubing into diluent- wash sol. flask again.

10.3	Make sure waste lines tubes are placed in container.

10.4	Latch platen plate to compress manifold tubing by engaging the hooks with the holding pins.  Raise the cover plate away from the platen until it is fully open, then press the cover plate down until the latches lock.  Then lower the cover plate onto the top of the platen.  Set the power switch at the side of the pump to position on.  Set the red switch on top of the pump to position run and the black switch to position normal. The rollers will start rotating, and the status LED on top of the pump will light up in green.

10.5	Switch on the power to: computer, monitor, printer and analyzer modules (sampler, and colorimeter with cover on) and pump reagent water +wetting agent through reagents lines.

NOTE:	Sampler wash solution does not contain wetting agent.

10.6	When Windows 95 has started, click the START button in the lower left-hand corner of the screen and click AACE.  After starting AACE, the Main Menu screen will be displayed.  It contains the Menubar, Toolbar, System Window, Desktop and Taskbar.

10.7	Click on CONFIGURE menu to enter the configuration of your analyzing system and to     choose software options.  The entered configuration will influence the appearance of the menus when working with ACE.  Icons belonging to the system, e.g. the pump, the channels (colorimeter) and the sampler can be manually control by double-clicking them.

NOTE:	This is done in the first application to be prepared with AACE. You don’t need to enter the system configuration daily, unless, you change any setting.

10.8	Programming the Analysis.

10.8.1	Select SET UP-Analysis in the Main Menu or click the SET UP Analysis tool button. For each application to be run on your Analyzer, create a separate Analysis in Set-UP- Analysis, it will then be used as template for the runs. The list on the left shows the analyses(.anl) that are contained in the AACE/Data directory.  Each Analysis is stored in a separate subdirectory (folder).The list on the right shows the Analysis and run files(.run) belonging to the Analysis currently selected.  Each run consists of a number of files with different extensions.  *.run stores Main Page, Channel and Tray Protocol information.

10.8.2	Enter the following information:

10.8.2.1	Analysis Main Page, including general data, for example the sampling rate or sample and wash time.

| Name of Analysis | NH3-N & N03N02-N |
| --- | --- |
| Samples per Hour | 80 |
| Sample/Wash | 6.0 |
| Sample Time | 39 |
| Wash Time | 6.0 |
| Method1 | EPA 350.1 |
| Method 2 | EPA 353.1 |
| Note: A high sample to wash ratio gives more accurate results. The reason is that the steady-state plateau is increased, allowing a better average to be taken. |  |

10.8.2.2	Tray Protocol, reflects the arrangements of the cups on the sample tray.

On the left-hand side, the Tray Protocol shows a table with the following columns: peak, icon, type, cup, sample ID. The Primer must be the first cup on the tray. It starts the peak window and must reach a certain peak height . The highest standard (2.0ppm) is used as Primer. The Primer acceptance criteria are set in the Peak Detection tab in Configure- Software. The concentration of the standards must be entered for each channel. In a run, their peak height is measured, and sample concentrations are calculated by comparison of their peak height to that of the standards. The calibrants should come immediately after the Primer, in descending order of concentration (i.e., the highest standard first). On the right- hand side of the Tray Protocol are buttons to enter the cup types and their position on the sample tray. P, 7C,3N,CARRYOVER SET,2QCS, 10S,…2N... 2QCS,2N,E

10.8.2.3	Channel, containing data related to each channel, such as method name, calibrator concentrations and corrections.

| Units | Mg/L |
| --- | --- |
| Calibration Fit | Linear |
| Peak Window |  |
| Start | 50 |
| End | 105 |
| Autobase Reference (%) | 5 |
| Calibrants concentrations | 2.000, 1.000, 0.5000, 0.1000, 0.1000, 0.0500,0.0000 |
| Base Correction | NO |
| Drift Correction | NO |
| Carry Over Corr | YES |
| When setting up an Analysis or run, you can generate a report. This can be useful before preparing the cups on the sample tray. At the bottom of the window OK to store the settings. |  |

10.9	To select NH3-N&NO3/NO2-N Analysis, double-click the Data folder, then double-click it.  The Analysis and run files belonging to this Analysis will appear in the list on the right.  Then continue in one of the following ways:

10.9.1	Click NEW RUN to create a new run based on the Analyses template.

10.9.2	Select a run file in the list and click COPY RUN to create a new run based on the selected run.

# 11.0	PROCEDURE

11.1	CHARTING allows you to:

11.1.1	View and print the charts to check if the baseline is stable before starting a run.

11.1.2	Check the reagent absorbance (once a week or whenever you use new reagents or pump tubes, or when you start an Analysis for the first time). See Section 11.12.

11.1.3	Adjust baseline and sensitivity (once a week or whenever you use new reagents or pump tubes, or when you start an Analysis for the first time). See Section 11.13

11.1.4	Adjust the lamp intensity for AA3 Colorimeters.

11.1.5	Manually control the analyzer modules.

11.1.6	The cover of the Digital Colorimeter must be closed, because stray light could distort the dark current measurement which is carried out after the download.

11.2	To open Charting windows for the channels:

11.2.1	Click CHARTING in the System window on the AACE desktop, or

11.2.2	Click RUN in the Main Menu, then point to Chart.

The lower left corner of the system window displays the Analysis that was last used for Charting or a run.  When Charting is started, the Charting button in the System window will change to Print.  This button allows you to print the charts up to the current moment or save the corresponding Charting values to a file.  Make sure you print or save the charts before you stop Charting.  The Charting values are only temporarily stored and will be deleted when the charts are closed.

11.3	Click the right mouse button in a Charting window to access the following options:

11.3.1	Set Base: Select this option to automatically adjust the baseline reading to the value defined in the Run Chart tab in Configure- Software.  The default is 5%.

11.3.2	Set Gain: Select this option to automatically adjust the Gain value while the highest standard is in the flowcell.  The Gain value (95%) is adjusted to the Primer Height value entered in the Run Chart tab in Configure-Software.

11.4	Auto Lamp: Click this option to automatically adjust the intensity of the lamp in the AA3 colorimeter. When you adjust the lamp intensity you should check if the value on the scale considerably increases and/or the Reference value (see below) falls under 80%.  This can be an indication that the lamp needs to be replaced.  The values are shown as % of the 221 A/D converter resolution of the photometer.

11.5	To manually set Base and Gain values and lamp intensity, double-click the Channel icon for the desired colorimeter channel while Charting is running.  During Charting, each Charting window contains the following elements:

11.5.1	a scroll bar at the top to move forwards and backwards through the chart.

11.5.2	horizontal lines representing the percent full scale and vertical lines representing the minutes.

11.5.3	a blue line representing the chart reading over time.  The current value is displayed in the upper-right corner.

11.5.4	an indicator bar (bargraph) on the right.  The following colors are used:

yellow if the current reading is below the baseline level of 5%

green if the reading is between baseline and primer level of 95%.

red if the reading is above the primer level.

11.6	Check the water baseline; it should be stable and flat.  Check bubble pattern in all lines, especially flowcell waste line. If the bubbles are the wrong shape, in most cases there is insufficient wetting agent, or tubing is wrong sort or contaminated.  Check or adjust baseline, reagent absorbance and sensitivity when necessary.

11.7	Place the reagent lines from the wash receptacle into the reagent containers. Allow the reagent stream and bubble pattern to stabilize. To close the Charting windows and exit the Chart option, you have the following possibilities:

11.7.1	click Stop in the System window on the AACE desktop, or

11.7.2	click Run in the Main Menu, then point to Stop and click the system you want, or

11.7.3	click the Close button in each Charting window to stop each channel individually.

11.8	START A RUN.  With the system pumping reagents and the sampler tray loaded according to the tray protocol, observe the bubble pattern and make any other final checks that may be necessary before starting the run.  Click Run in the Main Menu, then point to Start and click the system you want, or in the System window on the desktop, click Run.  The appearing dialog box shows the Analysis folder currently selected on the left and the runs belonging to this Analysis on the right. Usually the software measures the baseline for 3 minutes or less for a stable baseline and calculates the deviation. The Status line displays message, for example “ Looking for stable baseline” or “Autobase in progress”.  If the baseline is stable, the actual run will start and the sample probe will move to the cups as specified in the tray protocol.  If the baseline stays noisy, the run will not start.After starting a run, a run chart will appear for each selected channel.  Each Run Chart window contains the following elements:

11.8.1	Window title, displaying the channel number, the system number and the method name.

11.8.2	Scroll bar at the top to move forwards and backwards through the chart.

11.8.3	Force Baseline tool button.

11.8.4	Calibration Curve tool button.  This tool button appears when the calibration has been completed.  If you click it, another window will open, showing the calibration curve for this channel.

11.8.5	Peak symbol. The peak symbol appears when the Primer has been found and the cups are being analyzed. The peak turn red during each time the peak window is active, i. e, when the peak-picking program looks for the actual peak. This time interval can be adjusted for each channel.

11.8.6	Result field, displaying the number, the type and the uncorrected concentration of the cup that was last analyzed.

11.8.7	Window pane with the chart.

The chart itself shows:

11.8.8	A grid of horizontal lines representing the percent full scale and vertical lines representing the minutes.

11.8.9	A blue line representing the chart reading over time.

11.8.10	Green markers in the chart showing the approximate place where the program takes a peak.

11.8.11	An indicator bar on the right if View Bar graph was activated during Charting.

11.9	When the chart completely fills the window, it will automatically move to the left.  A run will stop automatically when all the samples have been analyzed and the final baseline as soon as it is steady is taken.  The window with the curve will disappear, and a message saying that the run was successfully completed will appear instead.  Automatically a report corrected results for standards and samples, charts and calibration curves are printed out as selected in the Configure Software- Post Run.

11.10	The Retrieve Menu.  After the completion of a run you can view and print charts, calibration curves, reports, raw data of a run, using the Retrieve option in the Main Menu if necessary.

11.11	System Shutdown.  The following describes the steps that are required if the pump will be out of operation for more than an hour, e.g., overnight.

11.11.1	Place the reagent lines and the sampler probe into wash solution.

11.11.2	Set the pump to fast speed and pump wash solution through the complete system for a few minutes until the tubing is clean.

11.11.3	Set the red switch to position stop.

11.11.4	Remove the pump platen to relieve pressure on the pump tubes.

11.11.5	Set the power switch at the side of the pump to position off.

11.11.6	Switch off all AA3 modules.

11.11.7	To exit AACE, select File-Exit in the Main Menu.

11.11.8	To shut down the computer , click the START button in the lower left-hand corner of the screen, and then click SHUT DOWN.

11.12	Reagent Water check relative to Water.  Repeat this procedure for each channel.

11.12.1	With the system pumping Reagent water+wetting agent and Charting windows displayed on the screen, wait until the water baseline is stable and the bubble pattern is regular.

11.12.2	Double-click the Channel icon in the System window to open the Base&Gain dialog box and set the Gain value to 10.  This sets full scale absorbance to 1.

11.12.3	Click the right mouse button in the chart and select Auto lamp from the local menu to adjust the lamp intensity.  The Reference channel should reach 80%.  If not, the lamp need to be replaced.

11.12.4	Click the right mouse button in the chart and select Auto Base from the local menu. This will adjust the baseline to the Chart % entered in Configure-Software-Run Chart, to 5%.

11.12.5	Write down the baseline reading in %, as indicated in the chart.

11.12.6	Introduce reagents into the system and wait until they arrive in the flowcell.

11.12.7	When the reagent baseline is stable, write down the new reading.

11.12.8	Calculate the reagent absorbance as follows:

Reagent absorbance relative to water = Reagent baseline(%) − Water baseline(%)100%

11.12.9	Check that the calculated reagent absorbance relative to water is ±15% of the value specified in the method description.

| Reagent absorbance for NH3-N | = | 0.03 AU |
| --- | --- | --- |
| Reagent absorbance for N03N02-N | = | 0.02 AU |

11.12.10	If the calculated value exceeds this range, substitute each reagent and the wash solution to check for contamination.

11.12.11	Click OK to close and save the new settings or Cancel to close it and restore the old settings.

11.13	How to check Sensitivity. Repeat this procedure for each channel.

11.13.1	Make sure the system is pumping reagents, the baseline is stable and Charting windows are displayed on the screen.

11.13.2	Double-click the Channel icon in the System window and set the Gain value to 10.  This sets full scale absorbance to 1.

11.13.3	Click the right mouse button in the chart and select Auto Base from the local menu. This will adjust the baseline to the Chart % entered in Configure-Software-Run Chart, to 5%.  Write down the baseline reading in %, as indicated in the chart.

11.13.4	Aspirate a full-scale standard (should be a mixed standard if is used for more than a channel) into the sample probe.

11.13.5	Double-click the Sampler icon in the System window, enter the cup position containing the top standard (2.0 mg/L) under CUP Position and click Sample.  Wait until the standard reaches the flowcell, check for noise and drift.

11.13.6	Note signal value in % and calculate the Sensitivity (also called Absorbance)

Sensitivity = % Std. − %Baseline100%11.13.7	Check that the calculated sensitivity is ±15% of the value specified in the method description.

| Sensitivity at 2.0 mg/L NH3-N | = | 0.10 - 0.20 AU |
| --- | --- | --- |
| Sensitivity at 2.0 mg/L N03-N | = | 0.40 AU |

11.13.8	If the sensitivity is not correct, check is reagents are OK and if the heating bath temperature is correct.  For NH3 if chlorine reagent is OK. For NO3 if reduction stage is OK.

# 12.0	DATA ANALYSIS AND CALCULATIONS

12.1	Direct measurement of Nitrate, Nitrite Nitrogen is made using linear regression analysis.  Sample data is reported in mg/L.

12.2	If dilution was needed, multiply READING by DILUTION FACTOR.CTOR.

# 13.0	METHOD PERFORMANCE

13.1	Participates in USEPA performance evaluation studies.

13.2	Precision and Accuracy Quality Control Chart are done as required.

13.3	MDL is determined annually or when a new operator begins work.Using reagent water fortified at a concentration of two to three times the estimated instrument detection limit, take seven replicate and process through the entire analytical method.  Calculate the MDL as follows:

MDL = (SD) x (t)

| t | = | students’ t value for a 99% confidence level and a standard deviation estimate with n-1 degrees of freedom [t = 3.14 for seven replicates]. |
| --- | --- | --- |
| SD | = | standard deviation of the seven analyses. |

# 14.0	POLLUTION PREVENTION

14.1	The quantity of chemicals purchased should be based on expected usage during 	its shelf life.

14.2	Actual reagent preparation volumes should reflect anticipated usage and reagent stability.

14.3	All chemical wastes are collected until properly disposed of by a licensed hazardous waste contractor.

# 15.0	WASTE MANAGEMENT

15.1	Refer to the WES Environmental Management System (EMS) policy and SOPs regarding waste management.

# 16.0	REFERENCES

16.1	Methods for Chemical Analysis of Water and Wastewater, Rev 1978.

16.2	AutoAnalyzer 3-Bran+Luebbe.  Operation Manual & Methods.

# 17.0	TABLES AND VALIDATION DATA

| TABLE 1. Quality Control Elements and Acceptance Limits for EPA Method 353.1 Determination Of Nitrate, Nitrite Nitrogen Colorimetric- Automated, Hydrazine Reduction |  |  |  |
| --- | --- | --- | --- |
| QC Elements | Frequency | Acceptance Criteria | Corrective Action |
| Initial Calibration | Every sample batch | Correlation > 0.99 | Recalibrate |
| Calibration Blank | At the beginning, after Calibration, at a 10% frequency and at the END of the batch | < RDL | ###### Check laboratory process. |
| IPC | After calibration, at a 10% frequency and at the END of the batch | ± 10% Recovery | Reanalyze. If the second analysis is outside the limits, sample analysis must be discontinued. Recalibrate |
| Laboratory Reagent Blank | At least one with each batch | < RDL | Check laboratory process |
| Quality Control Sample (QCS) | After calibration and at the END of the run | ± 10% Recovery | ###### Performance of the determinative step of the method is unacceptable |
| Laboratory Duplicates | For a minimum of 10% samples or at least one with every analysis batch. |  |  |
| Laboratory Fortified Sample Matrix (LFM) | For a minimum of 10% samples or at least one with every analysis batch. At the same conc. used to prepare the LFB (should not be less than four times the MDL). | ± 10% Recovery | ###### Either matrix or solution related if the performance for that analyte is shown to be in control. If necessary, increase the amount used for spike and reanalyze. |
| Laboratory Fortified Blank (LFB) | One for batch, at the same conc. used to prepare the LFM. | ± 10% Recovery | Should be identified and resolved before continuing analyses. |
| Method Detection Limit (MDL) | Annually or a new operator. |  |  |

| TABLE 2. Accuracy of Nitrate Nitrite Nitrogen Analysis in Water Samples by EPA Method 353.1 |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Date | Based on Sample Numbers | Accuracy (% Recovery) |  |  |  |  |  |
|  |  | Mean | SDa | Warning Limits |  | Control Limits |  |
|  |  |  |  | Upper (UWL) | Lower (LWL) | Upper (UCL) | Lower (LCL) |
| 04/08/2003 | L2002112.11 to L2003022-3 | 95.7 | 7.34 | 110 | 81.0 | 118 | 73.7 |
| a SD = standard deviation. |  |  |  |  |  |  |  |

| TABLE 3. Precision of Nitrate Nitrite Nitrogen Analysis in Water Samples by EPA Method 353.1 |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Date | Based on Sample Numbers | Precision (RPD)b |  |  |  |  |  |
|  |  | Mean | SDa | Warning Limits ( 2 SD) |  | Control Limits ( 3 SD) |  |
|  |  |  |  | Upper (UWL) | Lower (LWL) | Upper (UCL) | Lower (LCL) |
| 04/08/2003 | L2002040-9 to L2003022-3 | 3.55 | 7.33 | 18.2 | 0 | 25.6 | 0 |
| a SD = Standard deviation b RPD = Relative percent difference |  |  |  |  |  |  |  |

| TABLE 4. Accuracy of Nitrate Nitrite Nitrogen Analysis in Reagent Water (LFB) by EPA Method 353.1 |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Date | Based on Sample Dates | Accuracy (% Recovery) |  |  |  |  |  |
|  |  | Mean | SDa | Warning Limits |  | Control Limits |  |
|  |  |  |  | Upper (UWL) | Lower (LWL) | Upper (UCL) | Lower (LCL) |
| 04/08/2003 | 04/18/2002 to 03/26/2003 | 93.3 | 4.95 | 103 | 83.4 | 108 | 78.5 |
| a SD = standard deviation. |  |  |  |  |  |  |  |

| TABLE 5. Method Detection Limit (MDL) for Nitrate Nitrite Nitrogen Analysis in Reagent Water by EPA Method 353.1 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Date of Study | No. of Samples Spiked (n) | Spiked Concentration (mg/L) | Accuracy (Mean % Recoverya) | Precision (SDb in mg/L) | MDLc (mg/L) |
| 04/14/2003 | 7 | 0.09 | 109 | 0.004 | 0.02 |
| a Recovery of spiked concentration b SD = standard deviation of mean concentration measured c Method Detection Level (MDL) is 0.02 mg/L. WES’s Reporting Detection Level (RDL) is 0.06 mg/L |  |  |  |  |  |