STANDARD OPERATING PROCEDURE

#### For

#### Mod METHOD of AEM:61-

#### DETERMINATION OF Bacteroidetes Human Specific

| SOP #: Bacteroidetes | REVISION #: 0 | DATE: March 2006 |
| --- | --- | --- |
|  | Page 1 of 14 |  |
| This and all other DEA/WES SOP and QA documents are available (read/print only) to all WES employees on the WES server (w:\dea-qap\SOPs & QA Docs). It should be noted that the controlled SOP & QA documents are only those viewed on-line on the WES server. If this is a printed copy, it is an uncontrolled version and may not be the latest version currently in use. |  |  |

#### Massachusetts Department of Environmental Protection

#### Division of Environmental Analysis

Senator William X. Wall Experiment Station

37 Shattuck Street

Lawrence, MA  01843

| Prepared by: |  |  |  |  |
| --- | --- | --- | --- | --- |
|  |  |  | Date: | March 9, 2006 |
| Approved by: | ###### Jean Tang, Environmental Analyst |  |  |  |
|  |  |  | Date: |  |
| Approved by: | ###### Ron Stoner, Microbiology Laboratory Supervisor |  |  |  |
|  |  |  | Date: | March 28, 2006 |
| Approved by: | ###### Robert Serabian, Quality Assurance Officer |  |  |  |
|  |  |  | Date: |  |
|  | ###### Oscar Pancorbo, Director |  |  |  |

TABLE OF CONTENTS

Page

LIST OF REVISIONS	3

LIST OF TABLES	4

1.0	SCOPE AND APPLICATION	5

2.0	SUMMARY OF METHOD	5

3.0	DEFINITIONS	5

4.0	INTERFERENCES	6

5.0	SAFETY	6

6.0	EQUIPMENT AND SUPPLIES	6

7.0	REAGENTS AND STANDARDS	8

8.0	SAMPLE COLLECTION, PRESERVATION AND STORAGE	9

9.0	QUALITY CONTROL	9

10.0	CALIBRATION AND STANDARDIZATION	10

11.0	PROCEDURE	10

12.0	DATA ANALYSIS AND CALCULATIONS	11

13.0	METHOD PERFORMANCE	11

14.0	POLLUTION PREVENTION	11

15.0	WASTE MANAGEMENT	11

16.0	REFERENCES	12

17.0	TABLES AND VALIDATION DATA	13

# LIST OF REVISIONS

| Rev. # | Date | Description of Revision | Page # |
| --- | --- | --- | --- |
| 0 | March 2006 |  |  |

Disclaimer:

References to trade names, commercial products and manufacturers in this SOP do not constitute endorsement.

# LIST OF TABLES

Page

TABLE 1.	Recommended filtration volumes for the fecal coliform membrane filtration test	13

TABLE 2.	14

# 1.0	SCOPE AND APPLICATION

1.1	This polymerase chain reaction (PCR) assay may be used to selectively detect human fecal sources in water and wastewater samples using two sets of human fecal source specific primers and a selected annealing temperature for PCR process.  This is a present/absent method for analyzing relatively large sample volumes assuming that the samples are not highly turbid.  Special caution should be taken when interpreting the results as there are known non-human fecal sources are prone to faults positive results. Work best when use with a combination of other source tracking methods.

# 2.0	SUMMARY OF METHOD

2.1	A well-mixed sample is drawn through a 0.4-m polycarbonate membrane filter which is then extract the DNA and the human specific markers are amplified, if exist, with PCR thermocycler; all PCR products then separated and observed with a UV fluorescent dye. Bands with specific molecular size are recoded as positive results.

# 3.0	DEFINITIONS

3.1	Lab Duplicates (LD1 AND LD2) -- Two aliquots of the same sample taken in the laboratory and analyzed separately with identical procedures. Analyses of LD1 and LD2 indicate the precision associated with laboratory procedures, but not with sample collection, preservation, or storage procedures.

3.2	Field Duplicates (FD1 AND FD2) -- Two separate samples collected at the same time and placed under identical circumstances and treated exactly the same throughout the field and laboratory procedures. Analyses of FD1 and FD2 give a measure of the precision associated with sample collection, preservation and storage, as well as with laboratory procedures.

3.3	Method Blank (MB) -- An aliquot of reagent water or other blank matrix that is treated exactly as a sample including exposure to all glassware, equipment, solvents, and reagents that are used with other samples. The MB is used to determine if method analytes or other interferences are present in the laboratory environment, the reagents, or the apparatus.

3.4	Field Reagent Blank (FRB) -- An aliquot of reagent water or other blank matrix that is placed in a sample container in the laboratory and treated as a sample in all respects, including shipment to the sampling site, exposure to sampling site conditions, storage, preservation, and all analytical procedures. The purpose of the FRB is to determine if method analytes or other interferences are present in the field environment.

3.5	Method Positive Control (MPC) -- An aliquot of reagent water or other blank matrix to which known quantities of the method analytes are added in the laboratory. The MPC is analyzed exactly like a sample, and its purpose is to determine whether the methodology is in control, and whether the laboratory is capable of making accurate and precise measurements.

3.6	Method Negative Control (MNC) -- An aliquot of reagent water or other blank matrix are added in the laboratory.  The MNC is analyzed exactly like a sample, and its purpose is to determine whether the methodology procedure is antiseptical.  The MNC is used to determine if method analytes or other interferences are present in the laboratory environment, the reagents, the apparatus, or contaminants introduced by analyst’s practice.

3.7	Stock Standard Solution (SSS) -- A concentrated solution containing one or more method analytes prepared in the laboratory using assayed reference materials or purchased from a reputable commercial supplier.

3.8	Deoxyribonucleic acid (DNA) – a common molecule which is usually found to encode genetic information about an organism.

3.9	Polymerase Chain Reaction (PCR) – a method for enzymatically and exponentially amplifying a selected region of DNA. The DNA fragment can be seen when stained on a gel or used in other experiments.

3.10	Primer

3.11	Human Marker

# 4.0	INTERFERENCES

4.1	The DNA debris found in reagents, supplies, and sampling or analytical apparatus will interfere with the PCR assay.

4.2	There are certain chemicals or cell products can interfere with PCR reaction. To prevent the interference, be sure that no alcohol, medium or agar when making PCR sample and reagent mixtures.

4.3	The accuracy of this PCR assay is highly endangered when samples are collected with known cat population.

# 5.0	SAFETY

5.1	The water and wastewater samples assayed for Bacteroidetes have the potential to be contaminated with septic material that may contain pathogenic microbes as well as toxic or carcinogenic compounds.  Certain reagents used in this assay or to make components of this assay such as bleach, antibiotics, enzymes, acids, bases, and dyes are irritating or toxic if ingested and/or exposed to the skin or mucosal membranes of the body.  Hence labcoats, gloves, cuffs, and safety glasses or a shield for eye protection should be worn when processing and assaying media and samples.

5.2	All personnel handling environmental samples known or suspected to contain or to have been in contact with human waste should be immunized against known disease causative agents.

5.3	It is the responsibility of the user(s) of this method to comply with relevant disposal and waste regulations.  Sharp objects must be disposed in “SHARPS” containers and objects or media exposed to samples must be autoclaved before washing or disposal.

# 6.0	EQUIPMENT AND SUPPLIES

6.1	Sample Bottles – One-liter wide-mouthed borosilicate glass or plastic.  Bottles must be chlorinated, thio buffered rinsed and sterilized by autoclaving.  A 1-L sample bottle must contain 5 mL of a 2% solution of sodium thiosulfate to eliminate the possible effect of left over chlorine.

6.2	Dilution Bottles – Sterile bottles of the appropriate size to prepare filter positive control samples to run with each sample batch.

6.3	Recording Autoclave (Gettinge PACS)

6.4	Filtration Units – The WES Microbiology Laboratory uses two stainless steel 6-funnel filtration manifolds and stainless steel as well as plastic filtration funnels.  Discard plastic funnels if they develop deep scratches on the inner surface.

6.4.1	Stainless steel funnel assemblies are wrapped in Kraft paper and placed on a stainless steel perforated rack for sterilization in the autoclave (15 minutes).  Immediately following sterilization, the stainless steel funnels are placed in the UV light box for storage.  The plastic funnels are wrapped in a clear plastic autoclavable instrument bag and autoclaved (15 minutes) prior to use.  Prior to use, open the wrap if any, the funnels are placed in the UV light box and exposed to UV for 15 minutes.

6.4.2	Filtration funnels are not reused until they are fully cleaned and follow the sterilization steps described previously.

6.4.3	The filtration manifold is connected to two glass flasks in series with an electric vacuum pump securing a pressure differential of 138 to 207 kPa. The flask closest to the vacuum pump is for the protection of the pump from carryover water and must not be allowed to fill with water.

6.5	Membrane Filters – Use 0.4-m pore size, 47-mm diameter pre-sterilized, polycarbonate membrane filters.  Use only those membrane filters that have been found, through adequate quality control testing and certification by the manufacturer, to exhibit, as compared to control membrane filters: 1) Full retention of the organisms to be cultivated; 2) Stability in use; 3) Freedom from chemical extractables that may interfere DNA extraction or inhibit PCR assay; and 4) Not change the results of PCR assay.  Preferably use fresh stocks of membrane filters, and if necessary, store them in an environment without extremes of temperature and humidity.

  - Forceps – Smooth flat forceps, without corrugations on the inner sides of the tips.  Sterilize before use by dipping in 95% ethyl or absolute methyl alcohol and then flaming.

6.7	Freezer, -70°C – Use a ultra freezer that maintains a temperature of -70 °C (must ≤.-65 °C).

6.8	Sterile 1.5-µL polypropylene centrifuge tubes, PCR-certified

6.9	Sterile PCR-certified pipette tips, 10-, 100-, and 1000-µL

6.10	Short-wavelength (254nm) UV box for sterilization

6.11	Bench-top ultra centrifuge (Micromax RF OM3590, International Equipment Company, Needham Heights, MA)

6.12	Thermocycler (Mastercycler, Eppendrof, Westbury, NY)

6.13	Power Supply (EC250-90, E-C Apparatus Co., Holbrook, NY)

6.14	UV transilluminator

6.15	Bench-top ultra centrifuge (Heraeus INS Magafuge 2.0R)

6.16	Camera and suitable software (Kodak 290)

6.17	Assorted Glassware including flasks, bottles, test tubes, graduated cylinders, and pipettes.

# 7.0	REAGENTS AND STANDARDS

  - Ultra filtered sterile Reagent-grade (ASTM Type I) water

  - Preparation of Buffered Rinse Water

    - To prepare the stock phosphate buffer solution, dissolve 34.0 grams of potassium dihydrogen phosphate (KH2PO4), in 500 mL of reagent-grade water, adjust to pH 7.2 ± 0.5 with 1 N NaOH and dilute to 1 L with reagent-grade water.  Store solution at 4C in a 1-L screw cap bottle.  Discard solution if it becomes turbid.

    - To prepare the magnesium chloride solution, dissolve 81.1 grams of magnesium chloride (MgCl2  6 H2O) in 1 L of reagent-grade water.  Store solution at 4C in a 1-L screw cap bottle.

    - In a large 5-gallon carboy, add 95 mL of magnesium chloride solution and 24 mL of phosphate buffer solution to 5 gallons of reagent-grade water.  Dispense the buffer from the carboy into 1-L bottles.  Place nine bottles of the buffer solution into the autoclave and autoclave at 121C for 45 minutes.

    - Remove the bottles of buffered rinse water from the autoclave and store at room temperature.

    - A sterility check must be performed for each batch of buffered rinse water prepared.  Add 50 mL of buffered rinse water and 50 mL of lauryl tryptose broth (2X) into a sterile 250-mL Erlenmeyer flask.  Incubate for 24 hours at 35C.  Discard any batch of water showing turbidity or growth after incubation.  Record the results of the sterility check in the autoclave log.

  - Preparation of 2% Sodium Thiosulfate Stock

    - Dissolve 10.0 grams of sodium thiosulfate (NaS2O3), in 500 mL of reagent-grade water in a 1-L screw cap bottle.

    - Place the buffer solution into the autoclave and autoclave at 121C for 30 minutes. (45 min if in contains 1 L in a 1.5 L bottle).

    - Remove the bottle from the autoclave and store at room temperature.

    - When use for rinsing bottle during cleaning process, add 20 ml to 8 L of reagent-grad water.

  - Preparation of 5‰ Sodium hypochloride cleaning solution

    - To prepare the 5‰ Sodium hypochloride cleaning solution, add 100 mL of commercially available bleach to 900 mL of reagent-grad water.

  - Preparation of 0.5M EDTA

    - Dissolve 18.61 g ethylenediaminetetraacetic acid (EDTA, disodium, dehydrate) to 100 mL of reagent-grad water in a beaker with stir

    - Add about 2 g of NaOH pellets while stirring to bring the pH to base (pH8.0).  Be careful to add the last few pellets slowly to avoid overshooting the pH.

    - Filter and autoclave the solution in a crew-cap bottle.

    - Remove the bottle from the autoclave and store at room temperature.

  - Preparation of guanidine isothiocynate (GITC) buffer

    - 5 M guanidine isothiocyanate

    - 100 mM EDTA – pH 8

    - 0.5% Sarkosyl

  - FastDNA kit (Bio 101, Lajolla, CA)

  - ExTaq PCR buffer (Takara Mirus Bio Co., Madison, WI)

  - ExTaq polymerase (Takara Mirus Bio Co., Madison, WI)

# 8.0	SAMPLE COLLECTION, PRESERVATION AND STORAGE

  - Samples must be aseptically collected in a bleached sterile disposable plastic bottle leaving at least 1” (2.5 cm) of headspace to allow for sufficient mixing of the sample prior to analysis.  Make sure that the appropriate concentration of sodium thiosulfate is added to avoid any left over bleach in the bottle (see Section 6.1)

  - Keep sample bottle closed until it is to be filled, fill the bottle without rinsing, replace cap immediately, and make sure the screw cap is tight.

  - Samples must be kept at  4°C from the time of collection to the time of analysis.

  - Non-potable water samples must be transported to the laboratory as quickly as possible, received at the laboratory no later than 6 hours from first sample collection, and analyzed within 2 hours of receipt at the laboratory – maximum holding time of 8 hours from collection.

  - All samples must be accompanied by the WES sample tracking/ chain-of-custody form filled out by the collector, including all applicable chain-of-custody signatures.

# 9.0	QUALITY CONTROL

9.1	Minimum quality control (QC) requirements for this method are initial demonstration of laboratory capability, and with each sample batch, analysis of method blanks, laboratory fortified blanks, method positive control samples, and QCS samples.

9.2	Method Blanks (MB) – Before processing any samples, the analyst must demonstrate that all glassware, equipment, and reagents are free of DNAs and other interferences.  Each time a set of samples is analyzed, an LRB must also be analyzed to demonstrate that DNAs and other interferences are not present in the analytical system.  If the LRB produces a positive result for PCR reaction, the analyst must determine the source of contamination and eliminate the interference before processing any samples.  The results of the batch that was processed with the LRB a positive result will be rejected.  The samples will need to be re-processed after clearance of the contamination.

9.3	PCR amplification – for each PCR batch: Including: 10% negative (at random location), 10% positive (SSS at random location), and 10% random lab duplicates for the samples set that is processed. Either positive, negative, or duplicate samples failed, the run is rejected. Redo the PCR amplification.

9.4	Gel electrophoresis – The products of PCR will be separated in sizes by 1% argarose gel together with a DNA size marker. A success gel has marker correctly separated, the bend of positive PCR product appears on correct location, according to the marker, and no bend show up in the well that no products were loaded. Re-run the gel if bends show up in the well that no products were loaded or if marker fails.

9.5	Image obtaining - Take gel image by Kodak 290 system and export images in tiff format. The image files are saved in server and backup every night.

# 10.0	CALIBRATION AND STANDARDIZATION

10.1	Refer to Laboratory Quality Assurance Plan for calibration and standardization procedures of laboratory equipment used for this analysis.

# 11.0	PROCEDURE

  - The sample volume filtered will depend on the expected Bacteroidetes density for the sample type to be tested (see Table 1).

11.1.1	The ideal sample volume is the maximum amount that can be filtered within 5 min. Filter three different volumes depending on the expected fecal coliform density.  When less than 20 mL is to be filtered, add approximately 10 mL of sterile buffered rinse water to the funnel before adding and filtering the sample volume, or add the sample volume into a sterile dilution bottle containing sterile buffered rinse water and then filter the entire bottle content.  This increase in water volume aids in uniform dispersion of the bacterial suspension over the entire effective membrane filter surface.  For volumes of 100-1000 mL, use a sterile graduated cylinder to measure the sample volume.  For sample volumes less than 100 mL, use a sterile wide-tip pipette.

Note:  Bacteria must not be suspended in dilution water for more than 20 minutes at room temperature as death or multiplication could occur.

  - Sterilize/Sanitize Filtration Units – Use sterile filtration units only of each sample as a minimum precaution to avoid accidental contamination.  After funnels being used, treat any further sample filtration as a new filtration series and sterilize all funnels and bases in use.

11.3	Label Petri Dishes – Label the bottom (side containing the agar) of each petri dish with the: 1) LIMS sample number; and 2) Dilution/volume filtered.

  - Filtration of Sample – Using sterile (flamed) forceps, place a sterile membrane filter (grid side up) over the porous plate of the filter base.  Attach funnel to the base and lock it in place.  Shake the sample vigorously at least 25 times and take sample volume.  Turn on the vacuum and filter the sample volume completely.  With filter still in place, rinse the interior surface of the funnel by filtering three 20- to 30-mL portions of sterile buffered rinse water.  Rinsing between samples prevents carryover contamination.  Turn off the vacuum, remove the funnel, fold the membrane quarterly and remove the membrane filter with sterile (flamed) forceps (touching only within the outer 1/8” of the filter), and place it on the petri dish.  Immediately place the Petri dish at -70°C for at least 15 min before further use.

  - Extract DNA

  - PCR: (Follow EPA guildeline to perform PCR. PCR steps are a modification from Berner & Field’s method (2000). The details of the procedure remain confidential until published.)

# 12.0	DATA ANALYSIS AND CALCULATIONS

12.1

# 13.0	METHOD PERFORMANCE

13.1	Precision and Accuracy Quality Charts are done as required

# 14.0	POLLUTION PREVENTION

14.1	The quantity of chemicals purchased is based on expected usage during its shelf life.

14.2	Actual reagent preparation volumes reflect anticipated usage and reagent stability.

14.3	Any dyes used for staining DNA samples in the gels to make DNA visible under UV light is collected and shipped as hazardous waste.

# 15.0	WASTE MANAGEMENT

15.1	For information, consult The Waste Management Manual for Laboratory Personnel.

15.2	WES laboratories fully comply with all applicable federal, state, and local environmental regulations.  All hazardous waste products from this and other methods are collected in the laboratory and then disposed of by a licensed hazardous waste management contractor.  WES is also committed to protecting the air, water, and land by minimizing and controlling all chemical releases from fume hoods, biological safety cabinets, and bench operations.  Refer to the WES Environmental Management System (EMS) policy and SOPs regarding waste management.

15.3	Dispose of all samples by autoclaving for a minimum of 30 minutes.  Plastic plates and tubes are separated from medium for recycling.

# 16.0	REFERENCES

16.1	Bernhard, A. E., and K. G. Field. 2000. Identification of nonpoint sources of fecal pollution in coastal waters by using host-specific 16S ribosomal DNA genetic markers from fecal anaerobes. Appl. Environ. Microbiol. 66:1587-1594.

16.2	Bernhard, A. E., and K. G. Field. 2000. A PCR assay to discriminate human and ruminant feces on the basis of host differences in Bacteroides-Prevotella genes encoding 16S rRNA. Appl. Environ. Microbiol. 66:4571-4.

# 17.0	TABLES AND VALIDATION DATA

TABLE 1. Recommended filtration volumes for the fecal coliform membrane filtration test. Sample Type; Recommended Filtration Volume (mL). 10000; 5000; 1000; 100; 10; 1; 0.1; 0.0001. Lakes, reservoirs; X; X. Wells, springs; X; X. Water supply intake; X; X; X. Natural bathing waters; X; X; X. Sewage treatment plant; X; X; X. Farm ponds, rivers; X; X; X. Stormwater runoff; X; X; X. Raw municipal sewage; X; X; X. Feedlot runoff; X; X; X. Sewage sludge; X; X; X.

| TABLE 2. |  |  |  |  |
| --- | --- | --- | --- | --- |