STANDARD OPERATING PROCEDURE

For

The Determination of PCB Toxic Congeners in Waters and Wastewaters

| SOP #: PCB Congeners in waters | REVISION #: 1.0 | DATE: February 2003 |
| --- | --- | --- |
|  | Page 1 of 17 |  |
| This and all other DEA/WES SOP and QA documents are available (read/print only) to all WES employees on the WES server (w:\dea-qap\SOPs & QA Docs). It should be noted that the controlled SOP & QA documents are only those viewed on-line on the WES server. If this is a printed copy, it is an uncontrolled version and may not be the latest version currently in use. |  |  |

Massachusetts Department of Environmental Protection

Division of Environmental Analysis

Senator William X. Wall Experiment Station

37 Shattuck Street

Lawrence, MA 01843

| Prepared by: |  |  | Date: | ###### December 12, 2001 |
| --- | --- | --- | --- | --- |
|  | ###### Pakorn Varanusupakul, Environmental Analyst |  |  |  |
| Revised by: |  |  | Date: | ###### February 20, 2003 |
|  | ###### Michael D. Bebirian |  |  |  |
| Approved by: |  |  | Date: |  |
|  | ###### Robert Serabian, Quality Assurance Officer |  |  |  |
| Approved by: |  |  | Date: |  |
|  | ###### Oscar Pancorbo, Director |  |  |  |

TABLE OF CONTENTS

Page

LIST OF REVISIONS	3

LIST OF TABLES	4

1.0	SCOPE AND APPLICATION	5

2.0	SUMMARY OF METHOD	6

3.0	DEFINITIONS	6

4.0	INTERFERENCES	7

5.0	SAFETY	8

6.0	APPARATUS, SUPPLIES AND EQUIPMENT	8

7.0	REAGENTS AND CONSUMABLE MATERIALS	9

8.0	SAMPLE COLLECTION, PRESERVATION, AND STORAGE	10

9.0	QUALITY CONTROL	10

10.0	CALIBRATION AND STANDARDIZATION	11

11.0	PROCEDURE	12

12.0	DATA ANALYSIS AND CALCULATIONS	13

13.0	METHOD PERFORMANCE	13

14.0	POLLUTION PREVENTION	13

15.0	WASTE MANAGEMENT	14

16.0	REFERENCES	14

17.0	TABLES AND VALIDATION DATA	15

# LIST OF REVISIONS

| Rev. # | Date | Description of Revision | Page # |
| --- | --- | --- | --- |
| 0 | December 2001 | None |  |
| 1.0 | February 2003 | Table 3, QC Table, added | 17 |

# LIST OF TABLES

Page

TABLE 1.	Retention times of PCB Toxic Congeners in Waters and Wastewaters	15

TABLE 2.	Method Detection Limits (g/L) for PCB Toxic Congeners in Waters and Wastewaters	16

TABLE 3.	Quality Control Elements and Acceptance Limits for the Determination of PCB Congeners in Water by Modified 608	17

# 1.0	SCOPE AND APPLICATION

1.1	This is a gas chromatographic-electron capture detector (GC-ECD) method applicable to the determination of coplanar toxic congeners that exhibit dioxin-like toxicity from waters and wastewaters. The congeners are listed according to a numbering system developed by Ballshmiter and Zell (BZ#). The following congeners can be determined using this method:

| ANALYTE | CAS # |
| --- | --- |
| 3,4,4’,5-tetrachlorobiphenyl (BZ# 81) | 70362-50-4 |
| 3,3’,4,4’-tetrachlorobiphenyl (BZ# 77) | 32598-13-3 |
| 2’,3,4,4’,5-pentachlorobiphenyl (BZ# 123) | 65510-44-3 |
| 2,3’,4,4’,5-pentachlorobiphenyl (BZ# 118) | 31508-00-6 |
| 2,3,4,4’,5-pentachlorobiphenyl (BZ# 114) | 74472-37-0 |
| 2,3,3’,4,4’-pentachlorobiphenyl (BZ# 105) | 32598-14-4 |
| 3,3’,4,4’,5-pentachlorobiphenyl (BZ# 126) | 57465-28-8 |
| 2,3’,4,4’,5,5’-hexachlorobiphenyl (BZ# 167) | 52663-72-6 |
| 2,3,3’,4,4’,5-hexachlorobiphenyl (BZ# 156) | 38380-08-4 |
| 2,3,3’4,4’,5’-hexachlorobiphenyl (BZ# 157) | 69782-90-7 |
| 2,2’,3,4,4’,5,5’-heptachlorobiphenyl (BZ# 180) | 35065-29-3 |
| 3,3’,4,4’,5,5’-hexachlorobiphenyl (BZ# 169) | 32774-16-6 |
| 2,2’,3,3’,4,4’,5-heptachlorobiphenyl (BZ# 170) | 35065-30-6 |
| 2,3,3’,4,4’,5,5’-heptachlorobiphenyl (BZ# 189) | 39635-31-9 |

The extraction conditions of this method are comparable to USEPA Method 608, which measure the multicomponent constituents in commercial polychlorinated biphenyl (PCB) mixtures (Aroclors), chlorinated pesticides, toxaphene, and chlordane. The extract derived from this procedure may be analyzed for these constituents by using the GC conditions prescribed in either USEPA Method 608.

1.2	This method has been validated in a single laboratory and method detection limits (MDLs) have been determined for the analytes above. Observed detection limits may vary between waters, depending upon the nature of interferences in the sample matrix and the specific instrumentation used.

1.3	This method is restricted to use by or under the supervision of analysts experienced in the use of GC and in the interpretation of gas chromatograms.

1.4	When this method is used to analyze unfamiliar samples for any or all of the analytes above, analyte identifications must be confirmed by a different polarity capillary column.

# 2.0	SUMMARY OF METHOD

2.1	A measured volume of sample of 1 L is solvent extracted with methylene chloride by shaking in a separatory funnel. The methylene chloride extract is isolated, dried and concentrated to a volume of 10 mL after solvent substitution with hexane. The extract is quantitatively analyzed by gas chromatography-electron capture detector (GC-ECD).

# 3.0	DEFINITIONS

3.1	Surrogate analyte - A pure analyte(s), which is extremely unlikely to be found in any sample, and which is added to a sample aliquot in known amount(s) before extraction and is measured with the same procedures used to measure other sample components. The purpose of a surrogate analyte is to monitor method performance with each sample.

3.2	Laboratory duplicates (LD1 and LD2) - Two sample aliquots taken in the analytical laboratory and analyzed separately with identical procedures. Analyses of LD1 and LD2 give a measure of the precision associated with laboratory procedures, but not with sample collection, preservation, or storage procedures.

3.3	Field duplicates (FD1 and FD2) - Two separate samples collected at the same time and place under identical circumstances and exactly treated the same throughout field and laboratory procedures. Analyses of FD1 and FD2 give a measure of the precision associated with sample collection, preservation and storage, as well as with laboratory procedures.

3.4	Laboratory reagent blank (LRB) - An aliquot of reagent water that is treated exactly as a sample including exposure to all glassware, equipment, solvents, reagents, internal standards, and surrogates that are used with other samples. The LRB is used to determine if method analytes or other interferences are present in the laboratory environment, the reagents, or the apparatus.

3.5	Field reagent blank (FRB) - Reagent water placed in a sample container in the laboratory and treated as a sample in all respects, including exposure to sampling site conditions, storage, preservation and all analytical procedures. The purpose of the FRB is to determine if method analytes or other interferences are present in the field environment.

3.6	Laboratory fortified blank (LFB) - An aliquot of reagent water to which known quantities of the method analytes are added in the laboratory. The LFB is analyzed exactly like a sample, and its purpose is to determine whether the methodology is in control, and whether the laboratory is capable of making accurate and precise measurements at the required method detection limit.

3.7	Laboratory fortified sample matrix (LFM) - An aliquot of an environmental sample to which known quantities of the method analytes are added in the laboratory. The LFM is analyzed exactly like a sample, and its purpose is to determine whether the sample matrix contributes bias to the analytical results. The background concentrations of the analytes in the sample matrix must be determined in a separate aliquot and the measured values in the LFM corrected for background concentrations.

3.8	Stock standard solution - A concentrated solution containing a single certified standard that is a method analyte, or a concentrated solution of a single analyte prepared in the laboratory with an assayed reference compound. Stock standard solutions are used to prepare primary dilution standards.

3.9	Primary dilution standard solution - A solution of several analytes prepared in the laboratory from stock standard solutions and diluted as needed to prepare calibration solutions and other needed analyte solutions.

3.10	Calibration standard (CAL) – A solution prepared from the primary dilution standard solution and stock standard solutions of the internal standards and surrogate analytes. The CAL solutions are used to calibrate the instrument response with respect to analyte concentration.

3.11	Quality control sample (QCS) -- a sample matrix containing method analytes or a solution of method analytes in a water miscible solvent, which is used to fortify reagent water or environmental samples. The QCS is obtained from a source external to the laboratory, and is used to check laboratory performance with externally prepared test materials.

3.12	Method Detection Limit (MDL) -- The minimum concentration of an analyte that can be identified, measured and reported with 99% confidence that the analyte concentration is greater than zero.

3.13	External Standard (ES) -- A pure analyte(s) that is measured in an experiment separate from the experiment used to measure the analyte(s) in the sample. The signal observed for a known quantity of the external standard(s) is used to calibrate the instrument response for the corresponding analytes(s). The instrument response is used to calculate the concentrations of the analyte(s) in the sample.

# 4.0	INTERFERENCES

4.1	Method interferences may be caused by contaminants in solvents, reagents, glassware, and other sample processing apparatus that lead to discrete artifacts or elevated baselines in GC-ECD. All reagents and apparatus must be routinely demonstrated to be free from interferences under the conditions of the analysis by analyzing laboratory reagent blanks as described in Sect. 9.0.

4.1.1	Glassware must be scrupulously cleaned. Clean all glassware as soon as possible after use by thoroughly rinsing with the last solvent used in it. Follow by washing with hot water and detergent and thorough rinsing with tap and reagent water. Drain dry, and heat in an oven or muffle furnace at 400C for 1 hour. Do not heat volumetric ware. Thermally stable materials such as PCBs might not be eliminated by this treatment. Thorough rinsing with acetone may be substituted for the heating. After drying and cooling, seal and store glassware in a clean environment to prevent any accumulation of dust or other contaminants. Store inverted or capped with aluminum foil.

4.1.2	The use of high purity reagents and solvents helps to minimize interference problems. Purification of solvents by distillation in all-glass systems may be required.

4.1.3	WARNING: When a solvent is purified, stabilizers added by the manufacturer are removed, thus potentially making the solvent hazardous. Removal of preservatives by distillation may also reduce the shelf-life of the solvent.

4.2	Interferences by phthalate esters can pose a major problem in PCB toxic congeners analysis when using the electron capture detector. These compounds generally appear in the chromatogram as large peaks. Common flexible plastics contain varying amounts of phthalates that are easily extracted or leached during laboratory operations. Cross contamination of clean glassware routinely occurs when plastics are handled during extraction steps, especially when solvent-wetted surfaces are handled. Avoiding the use of plastics in the laboratory can minimize interferences from phthalates. Exhaustive cleanup of reagents and glassware may be required to eliminate background phthalate contamination.

4.3	Interfering contamination may occur when a sample containing low concentrations of analytes is analyzed immediately following a sample containing relatively high concentrations of analytes. Between-sample rinsing of the sample syringe and associated equipment with hexane can minimize sample cross contamination. After analysis of a sample containing high concentrations of analytes, one or more injections of hexane should be made to ensure that accurate values are obtained for the next sample.

4.4	Matrix interferences may be caused by contaminants that are co-extracted from the sample. The extent of matrix interferences will vary considerably from source to source, depending upon the water sampled. Cleanup of sample extracts may be necessary. Positive identifications should be confirmed (Section ).

4.5	WARNING: A dirty injection port liner will cause poor chromatography.

# 5.0	SAFETY

5.1	The toxicity and carcinogenicity of chemicals used in this method have not been precisely defined; each chemical should be treated as a potential health hazard, and exposure to these chemicals should be minimized. The laboratory is responsible for maintaining a current awareness file of Occupational Safety and Health Administration (OSHA) regulations regarding the safe handling of the chemicals specified in this method. A reference file of material safety data sheets (MSDSs) should also be made available to all personnel involved in the chemical analysis.

5.2	Pure standard materials and stock standard solutions of target compounds must be handled in a fume hood. A NIOSH/MESA approved toxic gas respirator should be worn when the analyst handles extremely high concentrations of these toxic compounds. Safety glasses and protective clothing should be worn. If skin or eye contact occurs, large volumes of water are applied to flush the area of contact. An emergency shower and eyewash station are located in the hallway immediately outside of the laboratory.

# 6.0	APPARATUS, SUPPLIES AND EQUIPMENT

6.1	Sample bottle -- Borosilicate, amber, 1-L volume with PTFE lined screw cap.

6.2	Glassware.

6.2.1	Separatory funnel -- 2000-mL, with PTFE stopcock, ground glass or PTFE stopper.

6.2.2	Flask, Erlenmeyer -- 500-mL.

6.2.3	Kuderna Danish (K-D) concentrator tube, 10- or 25-mL, graduated (Kontes K-570050-1025 or K-570050-2525 or equivalent). Calibration must be checked at the volumes employed in the test. Ground glass stoppers are used to prevent evaporation of extracts.

6.2.4	K-D evaporative flask, 500-mL, (Kontes K-570001-0500 or equivalent). Attach to concentrator tube with springs or clips.

6.2.5	Snyder column, three-ball (Kontes K-503000-0121 or equivalent).

6.2.6	Vial -- Glass, PTFE lined screw cap, 5- to 10-mL capacity.

6.2.7	Funnel – Glass, large

6.2.8	Cylinder – Various.

6.2.9	Pasture pipette – Disposable.

6.3	Boiling stones – Carborundum, #12 granules (Arthur H. Thomas Co. #1590-033 or equivalent). Heat at 400 C for 30 min prior to use. Cool and store in a desiccator.

6.4	Water bath -- Heated, capable of temperature control ( 2 C). The bath should be used in a hood.

6.5	Analytical balance -- Capable of accurately weighing to the nearest 0.0001 g.

6.6	Gas Chromatograph - Varian CP-3800 GC with dual ECD detectors equipped with Varian 8200, 48-sample tray autosampler system. The Varian 3800 system is capable of temperature programming and is equipped with electronic pressure control. The data system is a Varian Star Workstation that is operated with Windows 98 and a Dell OptiPlex GX1 350 Mhz Pentium II computer workstation.

6.6.1	Column 1 (Primary column): DB-5, 60 m long x 0.25 mm ID bonded fused silica column, 0.25 m film thickness.

6.6.2	Column 2 (Confirmation column): DB-1701, 60 m long x 0.25 mm ID bonded fused silica column, 0.25 m film thickness.

6.6.3	Helium carrier gas flow is established at 38 cm/sec linear velocity and 95%Argon/5%methane is set at 25 mL/min as make-up gas.

6.6.4	The oven temperature is programmed by holding the oven at 100C for 1 min then ramping to 200C at 15C/min and held for 30 min. The oven is then programmed to 280C at 5C/min and held for 10 min. Total run time is calculated to be 63.67 min. The injector temperature is set at 290C and both detector temperatures are held constant at 320C. The injection volume is 5 L split/splitless mode with a 1 min vent delay to achieve a 50:1 split ratio. After 10 min, a 5:1 split ratio is held constant to conserve carrier gas.

6.7	Detectors -- Electron capture detector (2-Ni63 high activity).

6.8	TABLE 1 lists retention times observed for method analytes using the columns and analytical conditions described above.

# 7.0	REAGENTS AND CONSUMABLE MATERIALS

7.1	Acetone, methylene chloride, hexane and methyl-tert-butyl ether (MTBE) -- Pesticide grade or equivalent.

7.2	Sodium Sulfate, granular, anhydrous, ACS grade. Heat treat in a shallow tray at 450 C for a minimum of 4 hours to remove interfering organic substances.

7.3	Glass Wool—Pesticide grade or solvent washed with hexane.

7.4	Disposable Pasteur pipettes - Borosilicate glass.

7.5	Reagent Water – ASTM Type I reagent water.

7.6	Standard solutions may be purchased from EM Science or Ultra Scientific as certified solutions. Purchased stock standard solutions should be replaced after the indicated expiration date has surpassed or, sooner if comparison with laboratory fortified blanks or QC samples indicate a problem.  Intermediate solutions should be replaced after six months or, sooner if comparison with instrument response factors indicate a problem.

7.6.1	PCB congeners (BZ#81, BZ#77, BZ#123, BZ#118, BZ#114, BZ#105, BZ#126, BZ#167, BZ#156, BZ#157, BZ#180, BZ#169, BZ#170, BZ#189), 100 g/mL in methanol.

7.6.2	Pentachloronitrobenzene (PCNB) solution, 100 g/mL in MTBE for use as surrogate standard.

7.6.3	Pesticides surrogate standard spiking solution, 200 g/mL in acetone (Decachlorobiphenyl (DCBP) and Tetrachloro-m-Xylene (TCMX))

7.7	Intermediate PCB congeners are prepared by accurately pipetting 100 L of each 100 g/mL PCB congener into a 10.00 mL volumetric flask partially filled with hexane. The intermediate is brought to the volume with hexane. The final concentration of the intermediate is 1.0 ng/L. The solution is transferred to a PTFE lined screw cap vial. Store in refrigerator at 4 C and protect from light.

7.8	Surrogate standard spiking solution is prepared by accurately pipetting 500 L of 200 g/mL pesticides surrogate standard spiking solution, and 1000 L of 100 g/mL of PCNB solution into a 50.00 mL volumetric flask partially filled with MTBE. The solution is brought to the volume with MTBE. The final concentration is 2.0 ng/L. Store the flask in the refrigerator at 4C and protect from light.

# 8.0	SAMPLE COLLECTION, PRESERVATION, AND STORAGE

8.1	Grab samples must be collected in glass containers.

8.2	Sample preservation - Samples must be refrigerated at 4 C from the time of collection until extraction. A sample holding time of 14 days has been established for the water samples.

8.3	Extract storage - Sample extracts should be stored at 4 C away from light. A 14-day maximum extract storage time is recommended.

# 9.0	QUALITY CONTROL

9.1	Minimum quality control (QC) requirements are initial demonstration of laboratory capability, determination of surrogate compound recoveries in each sample and blank, analysis of laboratory reagent blanks, laboratory fortified samples, laboratory fortified blanks, and QC samples.

9.2	Laboratory Reagent Blanks -- All extraction sets include a laboratory reagent blank consisting of Type I water and surrogates.

9.3	The analyst is permitted to modify GC columns, GC conditions, GC detectors, continuous extraction techniques, concentration techniques (i.e. evaporation techniques), and internal standards or surrogate compounds. Each time such method modifications are made, the analyst must repeat the procedures for MDL determinations.

9.4	Assessing surrogate recovery

9.4.1	When surrogate recovery from a sample or method blank is <60% or >140%, (1) check calculations to locate possible errors, (2) fortifying solutions for degradation, (3) contamination or other obvious abnormalities, and (4) instrument performance. If those steps do not reveal the cause of the problem, reanalyze the extract.

9.4.2	If a blank extract reanalysis fails the 60-140% recovery criterion, the problem must be identified and corrected before continuing.

9.4.3	If sample extract reanalysis meets the surrogate recovery criterion, report only data for the reanalyzed extract. If sample extract reanalysis continues to fail the surrogate recovery criterion, report all data for that sample as suspect.

9.5	Assessing laboratory performance – Laboratory fortified blank.

9.5.1	The laboratory must analyze at least one laboratory fortified blank (LFB) sample with every twenty samples or one per sample set (all samples extracted within a 24-hr period) whichever is greater. The fortified concentration of each analyte in the LFB should be 10 times. Calculate accuracy as percent recovery (Xi). If the recovery of any analyte falls outside the control limits, that analyte is judged out of control, and the source the problem should be identified and resolved before continuing analyses.

9.5.2	Until sufficient data become available from within laboratory, usually a minimum of results from 20 to 30 analyses, the laboratory should assess laboratory performance against control limits. When sufficient internal performance data becomes available, develop control limits from the mean percent recovery (x) and standard deviation (s) of the percent recovery. These data points are used to establish upper and lower control limits as follows:

UPPER CONTROL LIMIT = x + 3s

LOWER CONTROL LIMIT = x - 3s

After each five to ten new recovery measurements, new control limits should be calculated using only the most recent 20-30 data points.

9.5.3	The laboratory periodically documents and determines its detection limit capabilities for the analytes of interest. This procedure is usually done once a year.

9.5.4	The laboratory analyzes a QC sample purchased from Ultra Scientific or EM Science at for least 10% of samples. The congener recovery must be between 60-140%.

9.6	Assessing method performance – Laboratory fortified sample matrix.

9.6.1	The laboratory adds a known concentration to a minimum of 10% of routine samples, or one sample concentration per set, whichever is greater.

9.7	The laboratory may adopt additional quality control practices for use with this method. The specific practices that are most productive depend upon the needs of the laboratory and the nature of the samples. For example, field or laboratory duplicates may be analyzed to assess the precision of the environmental measurements or field reagent blanks may be used to assess contamination of samples under site conditions, transportation and storage.

# 10.0	CALIBRATION AND STANDARDIZATION

10.1	Establish GC operating parameters equivalent to those indicated in Section . The GC system must be calibrated using the external standard technique.

10.2	External standard calibration procedure

10.2.1	Prepare calibration standards at a minimum of three concentration levels for each congener of interest from intermediate solution and dilute with hexane to the volume in 10 mL volumetric flasks. The concentration levels of PCB congeners are 0.001, 0.005, 0.01, 0.05 and 0.10 ng/L. The lowest standard should represent analyte concentrations near, but above their respective MDLs. The other concentrations should correspond to the range of concentrations expected in the sample concentrates, or should define the working range of the detector. The calibration standards must bracket the analyte concentrations found in the sample extracts.

10.2.2	Calibration standards of pesticides mixture may be used for calibration standards of surrogate compounds (TCMX, PCNB and DCBP). (See EPA Method 508 (ref. ) or EPA Method 608 (ref. )).

10.2.3	Starting with the standard of lowest concentration, analyze each calibration standard and tabulate response (peak area) versus the concentration in the standard. The results can be used to prepare a calibration curve for each compound. Alternatively, if the ratio of response to concentration (calibration factor) is a constant over the working range (20% RSD or less), linearity through the origin can be assumed and the average ratio or calibration factor can be used in place of a calibration curve.

10.2.4	The working calibration curve or calibration factor is verified on each working day by the measurement of a minimum of two calibration check standards, one at the beginning and one at the end of the analysis day. These check standards should be at two different concentration levels to verify the calibration curve. For extended periods of analysis (greater than 8 hrs.), it is strongly recommended that check standards be interspersed with samples at regular intervals during the course of the analyses. If the response for any analyte varies from the predicted response by more than  20%, the test must be repeated using a fresh calibration standard. If the results still do not agree, generate a new calibration curve.

# 11.0	PROCEDURE

11.1	Extraction and sodium sulfate cleanup

11.1.1	Measure volume of sample to 1 L with graduated cylinder. Record the sample volume to the nearest 5 mL. Pour the entire sample into a 2-L separatory funnel. Fortify the sample with 50 L of the 2.0 ng/L surrogate standard solution.

11.1.2	Add 60 mL methylene chloride to the sample bottle, seal, and shake for 30 sec to rinse the inner walls. Transfer the solvent to the separatory funnel and extract the sample by vigorously shaking the funnel for 2 min with periodic venting to release excess pressure. Allow the organic layer to separate from the water phase for a minimum of 10 min. If the emulsion interface between layers is more than one third the volume of the solvent layer, the analyst must employ mechanical techniques to complete the phase separation. The optimum technique depends upon the sample, but may include stirring, filtration of the emulsion through glass wool, centrifugation, or other physical methods.

11.1.3	Assemble a K-D concentrator by attaching a 10-mL K-D concentrator tube to a 500-mL K-D evaporative flask with spring or clip. Add 1 to 2 clean boiling stones to the K-D unit.

11.1.4	Prepare a glass funnel plugged with glass wool at the bottom and filled with about 20-g anhydrous sodium sulfate.

11.1.5	Pre-wet the sodium sulfate with methylene chloride and carefully pass the extract (bottom layer) through the sodium sulfate funnel. Collect the extract in the K-D concentrator unit.

11.1.6	Repeat step  and  two more times. Combine all the extracts in the K-D unit.

11.1.7	Rinse the sodium sulfate with about 20-30 mL methylene chloride and combine in the K-D unit.

11.2	Extract concentration and solvent exchange

11.2.1	Attach a 3-ball Snyder column. Pre-wet the Snyder column by adding about 1 mL methylene chloride to the top. Place the K-D apparatus on a hot water bath, 65 to 70 C, so that the concentrator tube is partially immersed in the hot water, and the entire lower rounded surface of the flask is bathed with hot vapor. Adjust the vertical position of the apparatus and the water temperature as required to complete the concentration in 15 to 20 min. When the apparent volume of liquid reaches 2 mL, remove the K-D apparatus and allow it to drain and cool for at least 10 min.

11.2.2	Carefully add 50 mL of hexane to the top of the Snyder column to the flask. Place K-D apparatus on the water bath so that the concentrator tube is partially immersed in the hot water. When the apparent volume of liquid reaches 2 mL, remove the K-D from the bath and allow it to drain and cool.

11.2.3	Remove the Snyder column. Thoroughly wipe the water around the joint between K-D concentrator tube and evaporative flask, and carefully remove the flask. Adjust the final volume to 10.00 ml with hexane. If further concentration is needed, use an air-blow down apparatus or TurboVap.

11.2.4	Transfer extract to a 10-mL PTFE lined screw cap vial, store, and refrigerated at 4 C until analysis by GC-ECD.

11.3	GC-ECD

11.3.1	Section  summarizes the recommended operating conditions for the gas chromatograph. Included in Table 1 are retention times observed using this method. Other GC columns, chromatographic conditions, or detectors may be used if the QC requirements of this method are met.

11.3.2	Calibrate the system daily by running calibration standards.

11.3.3	Inject 5 L of the sample extract for congener analysis. Record the resulting peak size in area units.

11.3.4	If the response for the peak exceeds the working range of the system, dilute the extract and reanalyze.

11.4	Identification of analytes

11.4.1	Identify a sample component by comparison of its retention time to the retention time of a reference chromatogram. If the retention time of an unknown compound corresponds, within limits, to the retention time of a standard compound, and is its presence is confirmed on the secondary column, then identification is considered positive.

11.4.2	The width of the retention time window used to make identifications should be based upon measurements of actual retention time variations of standards over the course of a day. Three times the standard deviation of a retention time can be used to calculate a suggested window size for a compound. However, the experience of the analyst should weigh heavily in the interpretation of chromatograms.

# 12.0	DATA ANALYSIS AND CALCULATIONS

12.1	Calculate analyte concentrations in the sample from the response for the analyte using the calibration procedure described in the external cal procedure (Section ).

12.2	In the external standard calibration procedure, determination of the amount of material injected is made by the peak response of the known versus the unknown. The concentration (C) in the sample can be calculated from Equation 1.

C μg/L=A⋅ V(t)V(i)⋅V(s)			Equation 1

where:

A	=	Amount of material (ng)

V(t)	=	Volume of total extract (mL)

V(i)	=	Volume of injected extract (L)

V(s)	=	Volume of extracted sample (L)

# 13.0	METHOD PERFORMANCE

13.1	For precision and accuracy data, see Table 2

# 14.0	POLLUTION PREVENTION

14.1	The laboratory generates aqueous and solvent wastes form this procedure. Wastes are documented in a chemical inventory and are stored in a hazardous waste room awaiting transport and disposal.

14.2	For information about pollution prevention that may be applicable to laboratory operations, consult ref. .

# 15.0	WASTE MANAGEMENT

15.1	This method used methylene chloride and hexane as the extraction solvents. Waste solvents are discarded in waste containers. Once the waste containers reach capacity, they are transferred to a hazardous waste drum in the hazardous waste storage room in the WES basement. The hazardous waste drum is transported off the premises and disposed of by a licensed hazardous waste management contractor. The laboratory has implemented a chemical inventory database to track purchases of solvents and waste generated by use of these chemicals.

15.2	WES laboratories fully comply with all applicable federal, state, and local environmental regulations.  WES is also committed to protecting the air, water, and land by minimizing and controlling all chemical releases from fume hoods and bench operations. Refer to the WES Environmental Management System (EMS) policy and SOPs regarding waste management. For more information on waste management, consult ref. .

# 16.0	REFERENCES

16.1	US EPA Method 508

16.2	US EPA Method 608

16.3	AOAC Method 983.21

16.4	U.S. Environmental Protection Agency. 1995. Methods for the Determination of Organic Compounds in Drinking Water.

16.5	American Chemical Society (ACS).  Less is Better: Laboratory Chemical Management for Waste Reduction.  ACS, Department of Government Relations and Science Policy, Washington, DC.

16.6	American Chemical Society (ACS). The Waste Management Manual for Laboratory Personnel. ACS, Department of Government Relations and Science Policy, Washington, DC.

# 17.0	TABLES AND VALIDATION DATA

| TABLE 1. Retention times of PCB Toxic Congeners in Waters and Wastewaters |  |  |
| --- | --- | --- |
| Analyte | Retention Time (min)a |  |
|  | Primary | Confirmation |
| BZ# 81 | 29.96 | 37.35 |
| BZ# 77 | 31.45 | 40.05 |
| BZ# 123 | 35.61 | 41.25 |
| BZ# 118 | 36.06 | 41.80 |
| BZ# 114 | 38.16 | 42.85 |
| BZ# 105 | 40.51 | 45.16 |
| BZ# 126 | 44.19 | 48.13 |
| BZ# 167 | 45.96 | 48.64 |
| BZ# 156 | 47.63 | 50.30 |
| BZ# 157 | 48.01 | 50.74 |
| BZ# 180 | 48.81 | 50.87 |
| BZ# 169 | 50.07 | 52.88 |
| BZ# 170 | 50.57 | 52.75 |
| BZ# 189 | 52.15 | 54.32 |
| a Columns and analytical conditions are described in section . |  |  |

| TABLE 2. Method Detection Limits (g/L) for PCB Toxic Congeners in Waters and Wastewaters |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Analyte | Date of study | #Spiked Samples | Spiked Conc. (g/L) | Accuracy (%Recovery)a | Precision (SDb in g/L) | MDL (g/L) |
| BZ# 81 | 8/28/2001 | 7 | 0.050 | 104% | 0.0034 | 0.011 |
| BZ# 77 | 8/28/2001 | 7 | 0.050 | 100% | 0.0020 | 0.0062 |
| BZ# 123 | 8/28/2001 | 7 | 0.050 | 90% | 0.0029 | 0.0091 |
| BZ# 118 | 8/28/2001 | 7 | 0.050 | 89% | 0.0030 | 0.0094 |
| BZ# 114 | 8/28/2001 | 7 | 0.050 | 88% | 0.0029 | 0.0092 |
| BZ# 105 | 8/28/2001 | 7 | 0.050 | 85% | 0.0059 | 0.018 |
| BZ# 126 | 8/28/2001 | 7 | 0.050 | 93% | 0.0027 | 0.0086 |
| BZ# 167 | 8/28/2001 | 7 | 0.050 | 77% | 0.0027 | 0.0085 |
| BZ# 156 | 8/28/2001 | 7 | 0.050 | 75% | 0.0026 | 0.0082 |
| BZ# 157 | 8/28/2001 | 7 | 0.050 | 74% | 0.0025 | 0.0080 |
| BZ# 180 | 8/28/2001 | 7 | 0.050 | 72% | 0.0033 | 0.010 |
| BZ# 169 | 8/28/2001 | 7 | 0.050 | 74% | 0.0022 | 0.0069 |
| BZ# 170 | 8/28/2001 | 7 | 0.050 | 71% | 0.0031 | 0.010 |
| BZ# 189 | 8/28/2001 | 7 | 0.050 | 68% | 0.0021 | 0.0067 |
| a Average recovery of spiked concentration b SD = Standard deviation of mean concentration measured |  |  |  |  |  |  |

| TABLE 3. Quality Control Elements and Acceptance Limits for the Determination of PCB Congeners in Water by Modified 608 |  |  |  |
| --- | --- | --- | --- |
| QC Elements | Frequency | Acceptance Criteria | Corrective Action |
| Laboratory Extraction Blank | Every sample batch | Concentration < MDL | Reanalyze sample set |
| Laboratory Fortified Blank | Every sample batch | 60 – 140% Recovery | ###### Reanalyze sample set |
| Field Blanks | If available | Concentration < MDL | Reanalyze sample set |
| Laboratory Fortified Matrix | 5% of samples | 60-140% recovery | Qualify sample as “Matrix Effect” |
| Duplicates | 5% of samples | < 30% RPD | ###### Qualify data |
| Sample Holding Time |  | 14 days | Flag sample and analyze, or contact collector to obtain new sample |
| MDL | annually | Comparable MDLs from previous data | Re-evaluate method performance |
| Surrogate (PCNB) | Every sample | 60 – 140% Recovery | Re-analyze sample |
| Initial Calibration | Every sample batch | r2 > 0.99 for all target analytes | Recalibrate and reanalyze samples |
| Continuing Calibration | End of sample batch or each 12 hr. period | < 30% RSD | Recalibrate and reanalyze samples |
| Resolution | Every sample set | Baseline separation of congeners | Cut end of column, or clean injection port on GC |